| CU Medical Systems, Inc. |  |
| --- | --- |
| Address: 130-1, Donghwagongdan-ro, Munmak-eup, Wonju-si, Gangwon-do, Republic of Korea |  |
| TEL. (+82) 33-747-7657 |  |
| FAX: (+82) 33-747-7659 |  |
| E-mail: ra@cu911.com |  |
| Homepage: www.cu911.com |  |
| Device Name: | Semi-Automated External Defibrillators including Pads/Electrodes |
| Brand Name: | i-PAD |
| Model Name: | AED: CU-SP1, CU-SP1 PLUS Pads: CUA1007S, CUA1102S |
| Document Title: | Software Development File |
| Document Number: | CRD-0601-VV-SDF-01 |
| Manufacturer: | CU Medical Systems, Inc. 130-1, Donghwagongdan-ro, Munmak-eup, Wonju-si, Gangwon-do, Republic of Korea |
| EC Representative: | Medical Device Safety Service GmbH (MDSS) Schiffgraben 41, Hannover 30175 Germany |

|  | Dept. / Title | Name | Date | Signature |
| --- | --- | --- | --- | --- |
| Prepared by | Software Research Team / Principal Research Engineer | Jae Hong Chung | 2025-06-20 |  |
| Reviewed by | Regulatory Affairs Team / Principal Research Engineer | Jong Hyung Woo | 2025-06-20 |  |
| Approved by | R&D Center/ Senior Director | Tae Jong Noh | 2025-06-20 |  |

This document is owned by CU Medical Systems, Inc. contains confidential proprietary trade secret information and must not be copied. Only the recipient for the specific use for which it was uses the document and the information it contains. This document must be returned immediately upon the request by CU Medical Systems, Inc.

Revision History

| NO | Revision History | Revision No | Date |
| --- | --- | --- | --- |
| 1 | First Issuance | 0 | 2022-03-03 |
| 2 | Revision by audit of MDR 2017/745 | 1 | 2023-10-06 |
| 3 | Revision by audit of MDR 2017/745 | 2 | 2025-06-20 |
| 4 |  |  |  |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |
| 8 |  |  |  |
| 9 |  |  |  |
| 10 |  |  |  |
| 11 |  |  |  |
| 12 |  |  |  |
| 13 |  |  |  |
| 14 |  |  |  |
| 15 |  |  |  |

Table of Contents

1.	Introduction	4

1.1.	Reference Documents	4

1.2.	Software Safety Class Determination	4

2.	Software Development Plan	5

3.	Software Requirement Analysis	5

4.	Software Architectural & Detailed Design	6

5.	Software Unit implementation and verification	16

5.1.	Implement Software Unit	16

5.2.	Software Unit Verification	16

6.	Software integration and integration testing	21

7.	Software System Testing	27

8.	Software Release for utilization at a SYSTEM level	28

8.1.1.	Software released Versions	28

9.	Software Risk Management PROCESS	29

9.1.	Risk Analysis	29

9.2.	Risk Evaluation	29

9.3.	Risk Control	30

9.4.	Evaluation of overall residual risk	31

9.5.	FMEA Table	31

10.	Software Configuration Management	32

11.	Software Problem Resolution	32

Appendix 1. IEC 62304 Checklist	33

Appendix 2. Traceability Matrix	50

Introduction

This document summarizes the activities and deliverables of the SP1 Series Software Ver. 2.03 (CU-SP1 and CU-SP1 PLUS’ software) development lifecycle to demonstrate compliance with IEC 62304 and CE MDR requirement

Reference Documents

- Development Plan:

  - [CRD-0601-PP-SDP-01] Software Development Plan

- Specifications:

  - [CRD-0601-DI-SRS-01] Software Requirements Specification

  - [CRD-0601-DI-SDS-01] Software Design Specification

- Verification Reports:

  - [SUT] Software Unit Test Report

  - [SIT] Software Integration Test Report

  - [SST] Software System Test Report

- Procedures & Guidelines:

  - [CUP-731] Product Development Work Procedure

  - [CUP-737)Design Change Procedure

  - [CUI-731.001] Software Development Guidelines

  - [CUI-731.002] Software Release Guidelines

  - [CUI-731.004] Software Maintenance Guidelines

- Procedures & Guidelines:

  - [CRD-0601-VV-CSR-01] Cybersecurity Report

Software Safety Class Determination

The safety class of this software has been identified as Class C. The detailed risk analysis and basis for this classification are described in the [CRD-0601-DI-SRS-01] Software Requirements Specification.

Software Development Plan

Activities, schedules, responsibilities, and deliverables related to development, verification, and release were conducted in accordance with the Software Development Plan [CRD-0601-PP-SDP-01] and Software Development Guidelines[CUI-731.001].

Software Requirement Analysis

All requirements, including functional, performance, interface, security, and risk control measures, are defined in the Software Requirements Specification [CRD-0601-DI-SRS-01].

Additionally, based on the results of the threat modeling and risk analysis described in the Cybersecurity Report [CRD-0601-VV-CSR-01], specific security requirements for Access Control, Data Integrity, etc., were identified and included in the Software Requirements Specification ([CRD-0601-DI-SRS-01]).

Table 3-1 Software Requirements Specification List

| SRS ID | Description |
| --- | --- |
| FUN-REQ-01 | The AED must be capable of measuring and analyzing the patient’s electrocardiogram (ECG) when the pads are attached during device operation. |
| FUN-REQ-02 | The AED must be capable of delivering electrical energy to the patient according to the patient’s ECG condition. |
| FUN-REQ-03 | The AED must provide CPR guidance to the user when CPR is required following ECG analysis. |
| FUN-REQ-04 | The AED must be capable of real-time monitoring of the device status during use. |
| UI-REQ-01 | The AED must provide instructions regarding the pads status and attachment after device self-check during operation. |
| UI-REQ-02 | The AED must be able to inform the user of the current operational step of the device. |
| SFT-REQ-01 | The AED must be capable of internally discharging the charged high-voltage energy to ensure the safety of both the patient and the user. |
| SFT-REQ-02 | The AED must continuously monitor the voltage charged in the high-voltage capacitor (H.V. CAP.) to ensure it does not exceed the threshold value. |
| MNT-REQ-01 | When the AED's power button is pressed, the device must perform a Power On Self-Test (POST). |
| MNT-REQ-02 | The AED must be capable of performing Periodic Self-Tests (PST) to always maintain a Ready-For-Use status. |
| MNT-REQ-03 | Upon initial power input (battery insertion), the AED must verify that the device is functioning properly through a Battery Insertion Test (BIT). |
| MNT-REQ-04 | The AED must be able to extract the treatment history for a patient. |
| NFN-REQ-01 | If the AED becomes unavailable during use, it must be able to notify the user of the device’s status. |
| NFN-REQ-02 | If the AED is unavailable, it must provide guidance for performing emergency actions. |
| NFN-REQ-03 | In Administration Mode, the user must be able to change options related to CPR Guide settings. |
| NFN-REQ-04 | The AED must allow configuration of hardware-related and initial setup options. |
| NFN-REQ-05 | The AED must operate a low-power system to minimize battery consumption during long-term storage. |
| NFN-REQ-06 | The AED must be able to store usage-related data in internal memory. |

Software Architectural & Detailed Design

The overall software architecture, components, interfaces, SOUP identification, and detailed design of each unit are described in the Software Design Specification [CRD-0601-DI-SDS-01].

Table. Software Design Specification List

Table 4-1 Software Architectural Design List

| SAD | Title | SAD Description |
| --- | --- | --- |
| SAD-EA-01 | (Main ARM) ECG measurement | Acquires and filters ECG data for analysis |
| SAD-EA-02 | (Main ARM) ECG rhythm analysis | Analyzes ECG waveform and heart rate to determine if shock is required |
| SAD-EA-03 | (Main ARM) Electrocardiogram analysis | Initiates ECG analysis upon Pads attachment and guides user with results. |
| SAD-EA-04 | (Main ARM) ECG analysis algorithm meeting sensitivity criteria | 1.VF sensitivity > 90% 2.VT sensitivity > 75% * ECG Specificity Requirement (per IEC 60601-2-4 Clause 201.107) |
| SAD-EA-05 | (Main ARM) ECG analysis algorithm meeting specificity criteria | ECG Specificity > 95%. * ECG Specificity Requirement (per IEC 60601-2-4 Clause 201.107) |
| SAD-SE-01 | (Main ARM) Start of electrical energy charging | Begins energy charging immediately upon Pads attachment per patient mode. |
| SAD-SE-02 | (Main ARM) Electrical discharge following ECG analysis | Determines energy delivery based on ECG analysis and guides the user accordingly. |
| SAD-SE-03 | (Main ARM) Therapy board control | Controls Therapy B/D unit: - Sends commands per charge/discharge phase. - Verifies charge status from received data. - Triggers system error if Therapy B/D fails. |
| SAD-PD-01 | (Main ARM) Pads connector connection status check | Evaluates Pads connector ADC to determine connection status: - Adult Pads inserted / Pediatric Pads inserted / No Pads inserted |
| SAD-PD-02 | (Main ARM) Pads internal memory verification | - Analyzes internal Pads memory for usage history and expiration. - When expiration of the pads is detected, a command is sent to the UI to output a voice prompt (expiration notification). |
| SAD-PD-03 | (Main ARM) Determination of Pad Storage Condition | The pads storage condition is determined based on the results of PST and BIT, using the pad impedance ADC measurement values. |
| SAD-PD-04 | (Main ARM) Pad Attachment Monitoring | During operation mode, the pad attachment status is monitored by evaluating the pad impedance ADC measurement values in real time. |
| SAD-BM-01 | (Main ARM) Battery Level Classification | Classifies battery level as Full, Half, Quarter, Pre-Alarm, and Low |
| SAD-BM-02 | (Main ARM) Remaining battery capacity calculation | Calculates remaining battery capacity based on usage and storage history. |
| SAD-BM-03 | (Main ARM) Low battery state detection | - Detects low battery state using battery voltage ADC and estimated remaining capacity. - If a low battery state is detected in Operation Mode, up to 10 cycles (ECG analysis to CPR guide) are performed. |
| SAD-BM-04 | (Main ARM) Battery malfunction diagnosis | Judges battery malfunction based on voltage ADC indicating inoperable condition. |
| SAD-ST-01 | (Main ARM) POST (Power-On Self-Test) | Performs POST (Power-On Self-Test) upon entering Operation Mode. |
| SAD-ST-02 | (Main ARM) PST (Periodic Self-Test) | Performs PST (Periodic Self-Test) every 24 hours / weekly / monthly and logs results. |
| SAD-ST-03 | (Main ARM) BIT (Battery Insert Test) | Executes BIT (Battery Insert Test) upon battery installation. - non-interruptible. - The results of BIT shall be displayed on the device. |
| SAD-UI-01 | (Main ARM) UI Output Control (Beep, Audio, LED) | User Interface – Output: - Beep: Device stage & warning - LED: Status and stage indication - Audio: Guides via voice |
| SAD-UI-02 | (Main ARM) CPR guidance after energy discharge | After shock delivery, transitions to CPR Guide mode and provides voice guidance |
| SAD-UI-03 | (Main ARM) Pads application instructions | Guides user on correct Pads attachment. |
| SAD-UI-04 | (Main ARM) Device startup instructions | Announces device activation to the user. |
| SAD-UI-05 | (Main ARM) Patient mode determination | - Analyze the Adult/Pediatric Selection Switch and the inserted pad connector status to determine the patient mode, then notify the user accordingly. - If an abnormal value is entered during Adult/Pediatric mode selection, the default Adult Mode is passed to Pads Management |
| SAD-UI-06 | (Main ARM) CPR activity detection | Assesses CPR performance through CPR Guide behavior |
| SAD-UI-07 | (Main ARM) CPR guidance via voice output | Determines CPR Guide mode based on voice output setting. |
| SAD-UI-08 | (Main ARM) Pads application guidance via voice output | Executes Pads attachment guide based on configuration if Pads not attached. |
| SAD-UI-09 | (Main ARM) UI Input Handling (Button) | User Interface – Input: - 3 Button: Used for operation and selection - 1 Switch: Used for patient mode selection - 1 Microphone: Used to measure ambient noise so that the system can adjust the voice output environment. |
| SAD-DM-01 | (Main ARM) Norflash memory data storage | Stores system status, treatment history, configuration, and self-test results in Norflash memory. |
| SAD-DM-02 | (Main ARM) ECG data storage | Stores filtered ECG signals from the patient |
| SAD-DM-03 | (Main ARM) Extraction of therapy records | Exports stored treatment history externally. |
| SAD-DV-01 | (Main ARM) Wake-up state verification and initialization | Upon power input, verifies Wakeup State and initializes operational variables |
| SAD-DV-02 | (Main ARM) Power-off procedure (Main ARM Only) | Upon receiving a power-off signal, the system executes a shutdown routine—including EEPROM and ECG data storage, internal discharge, PST configuration, LCD/LED updates, and task cleanup—before issuing a final power-off command to the PIC. |
| SAD-DV-03 | (Main ARM) Error handling and categorization | - When a system error occurs, the error is detected, and the corresponding protocol is executed. The type of error is classified. - Upon receipt of an error code, the format and range are validated, and if an error is detected during reception or processing, re-reception or re-verification is attempted 2–3 times. |
| SAD-DV-04 | (Main ARM) Volume adjustment based on ambient noise | Analyzes ambient noise to adjust guide voice volume accordingly. |
| SAD-DV-05 | (Main ARM) CPR guidance in case of system error | Determines entry into Error CPR Guide mode if patient treatment is not possible |
| SAD-DV-06 | (Main ARM) State determination between wake-up and active modes | Evaluates Wakeup and button states from Main PIC to determine Active State. |
| SAD-DV-07 | (Main ARM) Administrator mode entry and subsequent processes | Administrator mode is accessed via a dedicated method, leading to therapy export or CPR setting mode. |
| SAD-DV-08 | (Main ARM) CPR guidance configuration updates | In CPR guide setting mode, confirms and modifies CPR guide configurations. |
| SAD-DV-09 | (Main ARM) Factory mode – Device setup | - In Factory Mode, performs memory initialization and HW calibration via device configuration entry. - Performs an integrity check on the device configuration data. [Target Data]: Voice Pack (File format verification) - Predefined default setting values are automatically applied to maintain normal operation. |
| SAD-DV-10 | (Main ARM) Factory mode – SD card-based device setup | - In Factory Mode using SD Card, modifies configuration and checks pad impedance ADC via voice feedback. - An integrity check is performed on the device configuration data. - Predefined default setting values are automatically applied to maintain normal operation. |
| SAD-DV-11 | (Main ARM) Factory mode – IrDA-based device configuration | In Factory Mode using IrDA, updates serial number, volume, and RTC time. |
| SAD-DV-12 | (Main ARM) Post-start device usage monitoring | If no user input is detected for 10 minutes after power is applied, the device determines that there is no intent to use. When the device determines that there is no intent to use, it stores the current state in internal NOR flash memory and then performs an automatic power-off procedure. This prevents unnecessary power consumption and blocks unintended operation by the user. |
| SAD-CA-01 | (Main ARM) SD card read/write operations | Uses SD Card as external memory for storing and retrieving data. |
| SAD-CA-02 | (Main ARM) Communication with Therapy PIC for high-voltage control | Communicates with Therapy PIC via UART to control high voltage modules. |
| SAD-CA-03 | (Main ARM) Communication with Main PIC for power and EEPROM control | Communicates with Main PIC via I²C for power management and EEPROM control |
| SAD-CA-04 | (Main ARM) IrDA communication | Controls IrDA communication via UART for device event transmission and setting data reception |
| SAD-CA-05 | (Main ARM) Data security management | Ensures integrity and confidentiality of data extracted/input from the device. - e.g., via AES-128, Hash, CRC |
| SAD-CM-01 | (Main PIC) Communication with Main ARM | Receives commands from Main ARM via I²C and transmits responses. |
| SAD-UM-01 | (Main PIC) UI Input Handling (Button) | User Interface – Input: - Button: Accepts user input for device operation |
| SAD-UM-02 | (Main PIC) UI Output Control (LCD, LED) | User Interface – Output(LCD, LED:): - LED: Error status, Pads connection status - LCD: Device, Pads, Battery indicators - Heartbeat ERROR LED: If the heartbeat signal from the main ARM is not detected, the red LED is forced to blink. |
| SAD-EE-01 | (Main PIC) EEPROM control | Controls EEPROM of Pads and battery (read/write). |
| SAD-PM-01 | (Main PIC) Main ARM activation | - Supplies power to Main ARM and determines Wakeup State before activation. - Upon receiving a power control command, the command code and bit values are validated to ensure correct interpretation. If the values are out of the allowed range or undefined, they are set to default safe values to prevent abnormal power shutdown. |
| SAD-PM-02 | (Main PIC) Main ARM status monitoring | Monitors Main ARM status based on output signal when power is supplied. |
| SAD-PM-03 | (Main PIC) Main ARM deactivation | - Cuts power to Main ARM upon shutdown command. - Upon receiving a Power OFF command, the command code and bit values are validated to ensure correct interpretation and to prevent power shutdown due to undefined values. |
| SAD-PM-04 | (Main PIC) Transition from Standby mode to a low-power state. | - Enters a low-power state if Main ARM power is disabled. |
| SAD-PM-05 | (Main PIC) RTC (Real-Time Clock) control | Reads RTC and measures standby time |
| SAD-CT-01 | (Therapy PIC) Communication with Main ARM | - Communicate with Main ARM via UART to receive commands and send responses. - The integrity of sensed values is verified through a dual-validation logic. If the sensed values are outside the allowed range or exhibit abnormal changes, charging and discharging control is immediately disabled. |
| SAD-HV-01 | (Therapy PIC) High voltage module control for Shock Energy charging | Controls high voltage module for energy charging: - Follows threshold settings for energy levels - Initiates/stops charge per Main ARM command - Verifies module state and charging progress |
| SAD-HV-02 | (Therapy PIC) External Discharge of Shock Energy | Delivers stored shock energy externally via defibrillation pads. - Starts external discharge per Main ARM command - Verifies module state and discharging progress - The integrity of sensed values is verified through a dual-validation logic. If the sensed values are outside the allowed range or exhibit abnormal changes, charging and discharging control is immediately disabled. |
| SAD-HV-03 | (Therapy PIC) Internal electrical energy discharge | The device performs a function that safely discharges high-voltage charging energy through an internal resistance. This internal discharge is automatically executed under the following conditions: - When the ECG analysis result is determined to be non-shockable - When the determination changes from shockable to non-shockable - When the user does not press the Shock button within 15 seconds - When a power-off event occurs during device operation - When pad attachment is removed or the connector is disconnected while in a charged state [Additional] This function shall comply with IEC 60601-2-4, Clause 201.12.4.103 (requirements for the internal discharge circuit), and shall safely dissipate energy within the internal circuitry without delivering current externally. |
| SAD-HV-04 | (Therapy PIC) Shock Energy Delivery Waveform | Applies biphasic waveform when delivering shock externally |

Table 4-2 Software Design Specification List

| SDS ID | Definition |
| --- | --- |
| SDS-MA-001 | Upon entering operational mode, if there is no intended user input for 10 minutes, the Device stores the current Status in Norflash Memory and automatically shuts down. The user's intended input confirms whether the Device is attached to the patient's body when intended use begins. |
| SDS-MA-002 | When power is applied to the Main ARM, the Main ARM initializes temporary variables, including Device operation values stored in the Buffer, and loads Device configuration values and reference values stored in Norflash Memory into temporary variables. |
| SDS-MA-003 | When Main PIC activates the power supply for Main ARM, Main ARM verifies the Wakeup State information from Main PIC. Based on the analysis of Wakeup State and button input, it determines the Active State and invokes the corresponding Task. |
| SDS-MA-004 | In the Status where ARM's driving power is supplied, the Device continuously performs real-time data saving to external/internal memory, and upon detection of Power OFF, it sets the PST operation time based on whether there is a System Error or not. |
| SDS-MA-005 | The device usage history data includes all events, progress status, and self-test results that occurred during device operation. |
| SDS-MA-006 | The treatment history data stores filtered ECG data along with events occurring in the operational mode and the time of event occurrence. |
| SDS-MA-007 | Configuration Data includes Device Configured Values, Device Reference Values, and Passwords related to Device operation, and can be modified and saved in CPR configuration mode, Factory Mode, and Field Mode. |
| SDS-MA-008 | If malfunctions or system halts occur in an embedded system, the system detects the occurring phenomenon and initiates a reset. Upon resetting the system, the Device restarts from Operation Mode initialization. |
| SDS-MA-009 | Upon occurrence of System Error Status, initiate internal discharge and disconnect power to the High Voltage Control module. Display Device Status to enable users and administrators to classify System Error Status. |
| SDS-MA-010 | When a system error is detected in the Device, it recognizes the system error situation and classifies the Device as having a System Error Status. It analyzes the occurring situation to classify the type of System Error through error codes for further identification. The System Error codes are stored in the Device Status Information of the NorFlash Memory. |
| SDS-MA-011 | Based on the ADC measurement values classified into categories for adult connector insertion, pediatric connector insertion, and others, determine whether defibrillation pads connector insertion is present or not. |
| SDS-MA-012 | Retrieve the first usage record of PADS EEPROM data from the initial PIC received data to determine if it has been used. |
| SDS-MA-013 | Retrieve the expiration date data of PADS EEPROM data from the main PIC received data to determine the expiration date of defibrillation pads. |
| SDS-MA-014 | The device determines whether the pad impedance ADC values measured during storage fall within the acceptable range for pad storage status. |
| SDS-MA-015 | In Operation Mode, if the measured pad impedance value falls within the reference range for pad attachment status, the system determines the pad as attached. |
| SDS-MA-016 | The defibrillation pad attachment guide mode can be determined using the attachment guide setting values. - Defibrillation Pad Placement Guide – Shortened Type: Provides brief voice instructions for pad placement intended for experienced users. - Defibrillation Pad Placement Guide – Detailed Type: Provides detailed voice instructions for pad placement intended for inexperienced users. |
| SDS-MA-017 | The remaining battery capacity (Capacity(%)) is determined through the usage history and storage history of the Device in the MAIN PIC received data, judged based on battery usage from the BATTERY EEPROM. |
| SDS-MA-018 | If the battery voltage ADC value is below the threshold value, it is determined as Battery Low Status. Battery Low Status indicates that the battery level is Low. |
| SDS-MA-019 | If the battery voltage ADC value falls below the threshold value, the device determines the status as "Device Unoperable due to Battery." The "Device Unoperable due to Battery" status assesses the battery level as "Depleted." |
| SDS-MA-020 | If the Battery Status is not Battery Low Status or Device Inoperable Due to Battery Status, classify the battery level according to the remaining battery capacity (Capacity(%)). |
| SDS-MA-021 | Measure ECG signal (ADC) values in Pads. |
| SDS-MA-022 | Remove unnecessary signal information to acquire ECG signals. |
| SDS-MA-023 | Store filtered ECG signals in internal memory (SRAM, NOR Flash). |
| SDS-MA-024 | Analyzes ECG rhythm characteristics to determine if defibrillation is necessary. |
| SDS-MA-025 | Analyzes the characteristics of ECG rhythms to determine the presence or absence of heartbeats. |
| SDS-MA-026 | When the measured value exceeds the threshold of the volume level measured at the micro level, audio output is set according to the measured value. - Maximum Volume Level: 10 (Maximum 90 dB) |
| SDS-MA-027 | Upon entering the CPR Guide, it determines whether to output the CPR Guide voice guide based on the user's i-Button input. |
| SDS-MA-028 | Upon entering the CPR Guide voice guidance mode, it determines the CPR Guide operation mode based on the settings for pad attachment guidance values and CPR Detailed Guide settings. |
| SDS-MA-029 | When the Status of the pads changes during CPR Guide operation, the Device recognizes the change in Status and proceeds by verifying the Status again after CPR Guide completion. |
| SDS-MA-030 | Based on the System Error detection value and battery low determination value, it determines whether to enter the Error CPR Guide. - Upon detecting a System Error during Operation mode and determining a low battery condition, proceed with the ERROR protocol and determine entry into the Error CPR Guide. |
| SDS-MA-031 | Analyze the ECG ADC values measured in the CPR Guide to determine whether the user is performing CPR on the patient. |
| SDS-MA-032 | If the Adult/Pediatric Selection Switch is set to Adult with the connector inserted for adult defibrillation pads, it is determined to be in adult mode. |
| SDS-MA-033 | Analyze the result value of the defibrillation pad connector connection status judgment and the Adult/Pediatric Selection Switch setting value to determine pediatric mode. - defibrillation pads connector connection status judgment result value has higher priority. - Determination of pediatric mode cases (1) Defibrillation pads connector inserted Status for pediatric use (2) When the Adult/Pediatric Selection Switch is set to Pediatric with the Adult defibrillation pads connector inserted, Status indicates Pediatric mode. |
| SDS-MA-034 | If there is a history of pad attachment, patient setup mode judgment based on Adult/Pediatric Selection Switch input shall not be performed unless only the defibrillation pad connector changes; otherwise, patient setup mode judgment cannot be altered. |
| SDS-MA-035 | If the Wakeup State in the received data of the Main PIC is BIT, proceed with BIT, or determine a separate Active State based on button combinations and entry passwords. - Entry Mode: BIT, Factory Mode, Field Mode, Device Configuration Mode determined via SD Card |
| SDS-MA-036 | When the Wakeup State in the received data of the Main PIC indicates PST, determine to enter the PST mode. |
| SDS-MA-037 | If the Wakeup State in the received data of the Main PIC indicates Normal, determine to enter the Operation Mode. |
| SDS-MA-038 | If the Wakeup State in the received data of the Main PIC indicates Information, determine to enter the administrator mode. |
| SDS-MA-039 | Upon successful completion of POST in Operation Mode, the device outputs the appropriate voice index regarding the connector connection status of the defibrillation pads, pad status, and current device operation status (patient setting mode). |
| SDS-MA-040 | In the defibrillation pad attachment guide of the operation mode, the voice prompt index of the defibrillation pad attachment guide is output based on the setting value of the defibrillation pad attachment guide. |
| SDS-MA-041 | During Pre-Charging in Operation Mode, based on the patient setting mode determination result value and adult mode electrical energy setting value, relevant commands for initiating charging are transmitted to Therapy B/D. |
| SDS-MA-042 | In the ECG Analysis Operation Mode, the measurement of ECG ADC values begins, and the Voice Prompt Index is output based on the ECG analysis results. |
| SDS-MA-043 | In the electrical energy discharge (internal/external) operation mode, discharge commands appropriate for each stage based on ECG waveform analysis results are stored in the Therapy PIC transmission data. Additionally, discharge commands are stored in the Therapy PIC transmission data based on ECG waveform analysis results, Shock Button input results, and pad attachment results during the electrical energy discharge phase. |
| SDS-MA-044 | The CPR Guide operation mode requests input of the CPR Guide i-Button upon entering the CPR Guide mode. Based on the CPR Guide i-Button recognition result value, proceed with the pause mode and voice guidance mode. Within the executed mode, the voice outputs the corresponding Voice Prompt Index according to the CPR Guide operational mode determination value, chest compression count setting value, breathing count setting value, cycle execution count value, chest compression speed setting value, pause duration setting value, ECG heartbeat analysis result value, and the specific CPR Guide progression. |
| SDS-MA-045 | The Administrator Mode can be entered from Standby Mode (with battery inserted status). In Administrator Mode, access to CPR Setup Mode and Treatment History Transmission Mode is available. Upon successful entry into Administrator Mode, the Device software version and recent usage history of the Device can be verified. |
| SDS-MA-046 | The Device Configuration Mode accessed via IrDA in Factory Mode can be entered using a separate method to input Device time and configuration data. Transmit time and configuration data from the Device to RTC and NorFlash Memory via IrDA on the PC. |
| SDS-MA-047 | The CPR guide setup mode can only be entered by administrators with appropriate privileges. In the CPR Guidance Setup Mode, you can check the current CPR Guidance Setup and modify the following items: - Items: Chest compressions Count, Breathing Count, Cycle Count, Chest Compression Rate, Pause Time, Detailed Guide Settings |
| SDS-MA-048 | Upon entering the treatment history transmission mode, it checks for the presence of an SD card. If an SD card is inserted, all stored treatment histories are transmitted to the SD card. If no SD card is inserted, one of the currently stored treatment histories is selected and transmitted via IrDA. |
| SDS-MA-049 | The Device settings mode of Factory Mode can be accessed through a dedicated entry method. |
| SDS-MA-050 | In the Device Configuration Mode of Factory Mode, default values can be stored in the Norflash Memory where Device Configuration values are stored by initializing the Norflash Memory. |
| SDS-MA-051 | In the Device configuration mode of Factory Mode, you can upload a Voice Pack to Norflash Memory. - Up to 10 languages can be uploaded. - The composition of the uploaded Voice Pack consists according to the product shipment priority, and Korean must be included as the first priority within the Voice Pack. |
| SDS-MA-052 | In the Device settings mode of Factory Mode, stored treatment history data can be initialized. |
| SDS-MA-053 | Upon entering the Device settings mode of Factory Mode, you can verify the version of the uploaded software on the Device up to the latest management version, and select/modify the desired settings item in Device settings mode. |
| SDS-MA-054 | In Device settings mode within Factory Mode, you can input reference values for Device operation and modify Device setting values. - Item Classification: ECG Reference Values, Impedance Reference Values, Electrical Energy Calibration Values, Adult Mode Electrical Energy Setting Values, Voice Variability Function Setting Values, - Detailed item (1) ECG Reference Values: ECG Gain, ECG Offset (2) Impedance Reference Value: 25 ohms Impedance Reference Value, 175 ohms (3) Electrical Energy Charging Threshold Calibration: Adult Energy Charging Threshold, Pediatric Energy Charging Threshold (4) Adult mode energy setting value: 150J (CU-SP1), 200J (CU-SP1 PLUS) (5) Voice modulation setting value: On/Off |
| SDS-MA-055 | In Device Setup Mode using an SD Card for Factory Mode, the requested settings specified at the factory specifications must be able to be changed simultaneously. |
| SDS-MA-056 | After configuring the device using an SD card, measure the impedance ADC value of the attached defibrillation pads and announce the result twice via voice guidance. |
| SDS-MA-057 | In Field Mode, you can change the multilingual settings and the settings for pad attachment guidance values. |
| SDS-MA-058 | Upon entering operational mode, the Device immediately checks if it is in normal Status. It verifies the Device Status and Device configuration values stored in Norflash Memory for any anomalies. It measures the battery voltage ADC value to determine if the battery is in low Status, then proceeds with charging and discharging tests to confirm if the Device Status allows for patient treatment. |
| SDS-MA-059 | Analyzes device inspection results to detect system errors and battery depletion. Records inspection results in system NORflash Memory and BATTERY EEPROM. Additionally, displays the status of inspection results on the device. |
| SDS-MA-060 | Daily PST performs a self-test on the Device's status every 24 hours since it was stored. During Daily PST, it checks the following: pad status, battery status, and comparison of device reference values and set values. |
| SDS-MA-061 | The Weekly PST performs an additional inspection for Device Status Self Test every seven days since the Device has been stored. After the Daily PST completes successfully, the Device additionally conducts an electrical charge test and battery status check. |
| SDS-MA-062 | The Monthly PST performs an additional inspection for High Voltage every 28 days since the Device has been stored. After Daily PST and Weekly PST are completed normally, the Device undergoes an additional inspection of the High Voltage system to confirm that it can be charged up to the set value. |
| SDS-MA-063 | Perform a Self Test according to the Device's storage date, analyze the Device inspection results to detect system errors, battery depletion, and pad faults. Record the inspection results in the system Norflash Memory and BATTERY EEPROM. If system errors, Battery Low Status, or pad fault conditions are detected, output an error alarm and display the corresponding Status on the Device. |
| SDS-MA-064 | Upon inserting a battery into the Device, a Self Test is initiated. During the Self Test performed upon battery insertion, the Device Status value in Norflash Memory (e.g., System Error) is reset, and a comprehensive inspection of all Device aspects is conducted to diagnose the current Device Status. BIT performs checks including pad status verification, battery status verification, verification of Device reference values and settings, button test, charging test, and battery status verification. |
| SDS-MA-065 | After inserting the battery, analyze the Device inspection results to detect system errors and insufficient battery levels. Record the inspection results in the system Norflash Memory and BATTERY EEPROM. Additionally, display the Status of the inspection results on the Device. |
| SDS-MA-066 | Based on the judgment of High Voltage Control commands according to the progression stages of operational modes and self-diagnostics, store them in the Therapy PIC transmission data. Based on the judgment result received by the Therapy PIC regarding the stage commands, store them in the Therapy PIC transmission data. |
| SDS-MA-067 | Upon receiving data (received data from Therapy PIC) analyzed after commands are transmitted to the Therapy PIC, it determines whether the response to the command is correct. If the response does not contain the requested data, it is indicated as a system error. If the response contains the requested data and transmits it normally, it is determined that the Therapy PIC command execution is complete, and proceed to the next step. |
| SDS-MA-068 | An ECG waveform classified as shockable is inputted to determine shockable status. (Waveform: Coarse VF, Fast VT) |
| SDS-MA-069 | Upon inputting a Non-Shockable ECG waveform, it determines as Non-Shockable (waveform: NSR, Asystole, ETC). |
| SDS-MA-070 | Communication transmit/receive data consists of temporary variables storing data scheduled for transmission before/after communication with external modules by the Main ARM. - Item: Therapy PIC Transmitted Data, Therapy PIC Received Data, Main PIC Transmitted Data, Main PIC Received Data, IrDA Transmitted Data, IrDA Received Data |
| SDS-MA-071 | The ADC measurement of defibrillation pads connectors measures the pad ID resistance within the connector, calculating the average of four measured ADC values as the ADC measurement value of the defibrillation pads connector. |
| SDS-MA-072 | The battery voltage ADC measurement involves measuring the resistance of the inserted battery voltage (12V) within the Device, considering the average of the four measured ADC values as the battery voltage ADC measurement value. |
| SDS-MA-073 | The Microphone ADC measurement determines the average value of 500 measured ADC values from the signal input through the microphone, considering this average as the Microphone ADC measurement value. |
| SDS-MA-074 | The temperature sensor ADC measurement considers the average of four measured ADC values obtained by measuring the temperature through the temperature sensor during device operation as the temperature sensor ADC measurement value. |
| SDS-MA-075 | Upon insertion of an SD Card, the Main ARM verifies the format of the inserted SD Card and reads the Device configuration file data from the SD Card. |
| SDS-MA-076 | Upon insertion of an SD Card, the Main ARM verifies the format of the inserted SD Card and copies the treatment record data stored in the Norflash Memory to the SD Card. |
| SDS-MA-077 | Upon data transfer from Therapy PIC to Main ARM, Main ARM immediately stores the received Therapy PIC data in temporary variables. |
| SDS-MA-078 | Upon storing data in the Therapy PIC transmit data, the Main ARM establishes communication with the Therapy PIC and transmits the stored data to the Therapy PIC. |
| SDS-MA-079 | The Main PIC stores the response from Main ARM in the Main PIC received data. |
| SDS-MA-080 | The Main ARM sends the Main PIC control command to the transmit data and then relays it to the Main PIC. |
| SDS-MA-081 | To transmit treatment history data via IrDA, copy the treatment history data into IrDA transmission data and transmit it to a PC-exclusive program via IrDA. |
| SDS-MA-082 | When receiving data using a dedicated program for PC on ARM, store it in the IrDA received data. |
| SDS-MA-083 | The Pads Impedance ADC measurement measures the Pads Impedance (the attached portion of Pads), calculating the Pads Impedance ADC measurement value as the average of four measured ADC values. |
| SDS-MA-084 | The ECG signal from the pad is measured using a 12-bit ADC with voltage resolution appropriate for ECG detection. |
| SDS-MA-085 | The RTC can be controlled to input time or read the current time. Set the RTC frequency to 32.768 Hz to verify the device time. |
| SDS-MA-086 | The LED indicates the status and progress stages of the Device. The default output value for each LED follows the hardware circuit design specifications. - LED for Device Status Verification: i-Button RED LED - LEDs for Device Status Awareness: Shock Button LED, Power Button LED, Do-Not-Touch-Patient Indicator, Pad Position Indicators, CPR Detection Indicator, i-Button BLUE LED |
| SDS-MA-087 | The beep sound output indicates Device progress stages and Device warnings. When indicating a device warning, emit a short beep sound. When indicating device progress stages, emit a long beep sound. |
| SDS-MA-088 | Output the voice index file determined by the progress stage through audio. Output the audio volume according to the set value. |
| SDS-MA-089 | Buttons recognize user input. Each button's default input value follows hardware circuit design specifications. - Button Item (Default Input Values per Button): i-Button (Low), Shock Button (High), Power Button (Low), Adult/Pediatric Selection Switch (High, High) |
| SDS-MA-090 | Analyze the ADC measurement values of the temperature sensor to determine if the storage temperature of the Device is cryogenic. |
| SDS-MA-091 | Encryption and decryption of data requiring encryption are performed using the AES-128 symmetric key algorithm. |
| SDS-MA-092 | A hash value is generated during file transfer to verify integrity. |
| SDS-MA-093 | Generate CRC values for packet data during data transmission. Utilize these values to verify the integrity of packet data. |
| SDS-MP-001 | When power is applied to Main ARM, Main ARM toggles the Status output on the PIC_TRIG pin every 10 ms. |
| SDS-MP-002 | The Main PIC continuously monitors the PIC_TRIG signal in real time. If the PIC_TRIG signal is not output, the Main PIC determines that the Main ARM is not operating normally. |
| SDS-MP-003 | If the self-diagnosis Count matches the PST operation time setting value, it transitions to the operating mode, and determines the Wakeup State as PST. |
| SDS-MP-004 | When the recognition result value for the Power Button input of Main PIC in standby mode is "Pressed", the system transitions to operational mode and determines the Wakeup State as Normal. |
| SDS-MP-005 | When the recognition result value for the input of the i-Button of the Main PIC in standby mode is "Cut", the system transitions to operation mode and determines the Wakeup State as Information. |
| SDS-MP-006 | If a System Dead Lock occurs in the Main PIC, determine that the Wakeup State is Reset. |
| SDS-MP-007 | Once the determination for Wakeup State in the Main PIC is completed, it assesses whether the battery level is within the operational range of the device. |
| SDS-MP-008 | Once the determination for Wakeup State in the Main PIC is completed, set the Power_ON_OFF pin (Main ARM power supply pin) to High and transition to the operational mode. - PWR_ON_OFF: Main ARM Power Supply Enable GPIO Pin |
| SDS-MP-009 | Upon executing the Power OFF protocol, set the Power_ON_OFF pin (main ARM power supply pin) to LOW and transition to standby mode. - PWR_ON_OFF: Main ARM Power Supply Enable GPIO Pin |
| SDS-MP-010 | When Main ARM disables power supply, it transitions to standby mode. In standby mode, Main PIC adjusts the system clock and reduces standby power to enable Main PIC to operate in low-power mode. |
| SDS-MP-011 | The RTC variable values consist of variables configured to read and set the counted time values from the data received by RCT. Item: Self-Diagnostic Count, RTC Data, PST Operation Time |
| SDS-MP-012 | The RTC's Clock output allows counting the passage of time from current time data. |
| SDS-MP-013 | Analyzes the received ARM data to determine the operation of the Main PIC. |
| SDS-MP-014 | Store the response according to the Main ARM command in the Main PIC transmitted data and then forward it to the Main ARM. |
| SDS-MP-015 | Communication transmit/receive data consists of temporary variables storing data scheduled for transmission or received prior to/after communication between the Main PIC and external modules. - Item: Main ARM Transmitted Data, Main ARM Received Data |
| SDS-MP-016 | The EEPROM communicates via 1-Wire communication and performs READ/WRITE operations on EEPROM data after issuing READ/WRITE commands following initialization. |
| SDS-MP-017 | If the PADS EEPROM responds with a Status signal, read the data from the PADS EEPROM and store it in the Device Status variable. |
| SDS-MP-018 | If the PADS EEPROM responds with a Status signal, record the PADS DATA stored in the Device Operational Status variable to the PADS EEPROM. |
| SDS-MP-019 | If the BATTERY EEPROM responds with a Status signal, read the data from the BATTERY EEPROM and store it in the Device Operation Status variable. |
| SDS-MP-020 | If the BATTERY EEPROM responds with a Status signal, record the BATTERY DATA stored in the Device operation Status variable to the BATTERY EEPROM. |
| SDS-MP-021 | The usage history of the patch shall be recorded in the patch EEPROM upon attachment of the patch. |
| SDS-MP-022 | Battery data recording shall be performed whenever the data of the relevant item is changed. - Item: Shock Count, BIT Count, PST Count, Battery Level, Usage Time |
| SDS-MP-023 | In standby mode, the Main PIC periodically sends signals to the PADS EEPROM to verify a response. If there is no response, determine the status as Defibrillation Pads Connector OFF, display the pad status on the LED and LCD, and record it in the temporary variable. If the previous response status and current status are not identical (e.g., Defibrillation pads connector OFF → Defibrillation pads connector ON), read the PADS EEPROM DATA into the buffer and display the pad status on the LED and LCD. |
| SDS-MP-024 | The RED status indicator indicates that the Device has an error status while in standby mode. - Error Status: Red LED blinks every 5 seconds |
| SDS-MP-025 | The green status indicator indicates whether defibrillation pads connectors are inserted when the device is in operational mode. - When defibrillation pads are not connected: Green LED blinks - When defibrillation pads connector is inserted: Green LED lights up |
| SDS-MP-026 | The default State of the Power Button remains LOW according to the HW design, and it is recognized as an input when held HIGH for more than 1 second. |
| SDS-MP-027 | The default state of the i-Button remains LOW according to the HW design, and it is recognized as an input when the button remains HIGH for more than 1 second. |
| SDS-MP-028 | The Device Operation of the Status LCD is indicated by O and X, displaying each item according to the system status. If the system status does not change, it maintains the displayed system status. - X: System Error occurs on the Device. Marked as X indicates that PADS and BATTERY on the Status LCD will also show Empty. - O: No System Error Occurs in Device |
| SDS-MP-029 | The PADS of the Status LCD display Full, Half, or Empty, according to commands from the Main ARM operation mode. During standby mode, the PADS of the Status LCD reflect responses from the PADS EEPROM signal and data from the PADS EEPROM. - If unused pads have more than 3 months remaining on their expiration date - Half: When unused pads have a shelf life remaining of 3 months or less - Defibrillation pads connector not inserted • Status expired or used pads • Expired or used pads |
| SDS-MP-030 | The BATTERY indicator on the Status LCD displays the battery level according to the battery status. |
| SDS-TP-001 | Set the electrical energy charging threshold value. Set the electrical energy charging threshold value corresponding to the received electrical energy value in Main ARM. - Condition: Charge threshold value can be set after mode change command - Set the critical value for electrical energy charging to 15 J. - Set the critical value for electrical energy charging to 50 J. - Set the critical value for electrical energy charging to 150J - Set the electrical energy charging threshold value to 200 J. |
| SDS-TP-002 | After ARM grants the charging command, the Therapy PIC proceeds with charging according to the charging threshold value set as a threshold value. |
| SDS-TP-003 | After issuing the charging stop command by ARM, the Therapy PIC verifies whether the operational status is in the "charging" status. If the status is charging, it proceeds with stopping electrical energy charging. |
| SDS-TP-004 | Apply electrical energy to Paz according to external discharge commands. |
| SDS-TP-005 | Upon receiving an internal discharge command, electrical energy is applied to the device's internal resistance. |
| SDS-TP-006 | Adjust the amount and timing of discharged electrical energy to output electrical energy in a biphasic waveform. |
| SDS-TP-007 | Capacitance of the capacitor must be measured using a High Voltage CAP ADC to monitor charging capacity. |
| SDS-TP-008 | Analyzes received ARM data to determine the operation of the Therapy PIC. |
| SDS-TP-009 | Store the response according to the Main ARM instruction in the Main ARM transmit data, then transmit it to the Main ARM. |

Software Unit implementation and verification

Implement Software Unit

The source code was written in compliance with the Software Development Guidelines [CUI-731.001] and is managed via the configuration management system.

Software Unit Verification

The verification activities and results for each unit are recorded in the Software Unit Test Report

Table. Software Unit Test Report List

| SUT-ID | Document Title |
| --- | --- |
| SUT-MA-001 | CRD-0601-0434-241217-001_PwrSleepTask |
| SUT-MA-002 | CRD-0601-0434-241217-002_InitRecord |
| SUT-MA-003 | CRD-0601-0434-241217-003__System_Read |
| SUT-MA-004 | CRD-0601-0434-241217-004_Dev_Operation_Data_Read |
| SUT-MA-005 | CRD-0601-0434-241217-005_Dev_Config_Data_Read |
| SUT-MA-006 | CRD-0601-0434-241217-006_AED_Main |
| SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
| SUT-MA-008 | CRD-0601-0434-241217-008_Power_Off |
| SUT-MA-009 | CRD-0601-0434-241217-009_PwrKeyProcTask |
| SUT-MA-010 | CRD-0601-0434-241217-010_setPSTIntervalCounter |
| SUT-MA-011 | CRD-0601-0434-241217-011_EventRecord |
| SUT-MA-012 | CRD-0601-0434-241217-012_ECGRecord |
| SUT-MA-013 | CRD-0601-0434-241217-013_Dev_Config_Data_UpDate |
| SUT-MA-014 | CRD-0601-0434-241217-014_ResetModeTask |
| SUT-MA-015 | CRD-0601-0434-241217-015_SystemMonitorTask |
| SUT-MA-016 | CRD-0601-0434-241217-016_SystemErrorMsg |
| SUT-MA-017 | CRD-0601-0434-241217-017_SystemErrorMsgduringCharging |
| SUT-MA-018 | CRD-0601-0434-241217-018_PadsConnectChkMsg |
| SUT-MA-019 | CRD-0601-0434-241217-019_PadsProcTask |
| SUT-MA-020 | CRD-0601-0434-241217-020_Pads_EEPROM_Use_Task |
| SUT-MA-021 | CRD-0601-0434-241217-021_Set_Pads_Exp_Flag |
| SUT-MA-022 | CRD-0601-0434-241217-022_PadsCheckTask |
| SUT-MA-023 | CRD-0601-0434-241217-022_PadsCheckTask |
| SUT-MA-024 | CRD-0601-0434-241217-024_RunningBatteryLevelCalculation |
| SUT-MA-025 | CRD-0601-0434-241217-025_BatteryLevelCalculation_withCounter |
| SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
| SUT-MA-027 | CRD-0601-0434-241217-027_StopSystemWhenBatteryDepleted |
| SUT-MA-028 | CRD-0601-0434-241217-028_BatteryMonitorTask |
| SUT-MA-029 | CRD-0601-0434-241217-029_ECGAnalSegProc |
| SUT-MA-030 | CRD-0601-0434-241217-030_ECGFilterOut |
| SUT-MA-031 | CRD-0601-0434-241217-031_ECGPreProcTask |
| SUT-MA-032 | CRD-0601-0434-241217-032_ECGAnalyzeTask |
| SUT-MA-033 | CRD-0601-0434-241218-001_AudioVolumeTask |
| SUT-MA-034 | CRD-0601-0434-241218-002_CPR_Selection_by_iBtn |
| SUT-MA-035 | CRD-0601-0434-241218-003_CPRProcTask |
| SUT-MA-036 | CRD-0601-0434-241218-004_CPRChgTask |
| SUT-MA-037 | CRD-0601-0434-241218-005_checkPadsStatusAfterCPR |
| SUT-MA-038 | CRD-0601-0434-241218-006_AEDPadsOff |
| SUT-MA-039 | CRD-0601-0434-241218-007_SysErrCPRChgTask |
| SUT-MA-040 | CRD-0601-0434-241218-008_SysErrCPRProtocolTask |
| SUT-MA-041 | CRD-0601-0434-241218-009_SysErrCPRProcTask |
| SUT-MA-042 | CRD-0601-0434-241218-010_CPRImpProcTask |
| SUT-MA-043 | CRD-0601-0434-241218-011_AdultPadsMsgTask |
| SUT-MA-044 | CRD-0601-0434-241218-012_InfantPadsMsgTask |
| SUT-MA-045 | CRD-0601-0434-241218-013_ModeChangeTask |
| SUT-MA-046 | CRD-0601-0434-241218-014_AED_mainStart |
| SUT-MA-047 | CRD-0601-0434-241218-015_FactoryModeChk |
| SUT-MA-048 | CRD-0601-0434-241218-016_FM_System_Field_Mode |
| SUT-MA-049 | CRD-0601-0434-241218-017_FM_SD_Card_Inserted |
| SUT-MA-050 | CRD-0601-0434-241218-018_BattInstOpTask |
| SUT-MA-051 | CRD-0601-0434-241218-019_EmergencyMsgTask |
| SUT-MA-052 | CRD-0601-0434-241218-020_ChargeProcTask |
| SUT-MA-053 | CRD-0601-0434-241218-021_AEDProcTask |
| SUT-MA-054 | CRD-0601-0434-241218-022_ShockProcTask |
| SUT-MA-055 | CRD-0601-0434-241218-023_NoShockProcTask |
| SUT-MA-056 | CRD-0601-0434-241218-024_Mute_CPR |
| SUT-MA-057 | CRD-0601-0434-241218-025_Mute_CPR_END |
| SUT-MA-058 | CRD-0601-0434-241218-026_Voice_CPR |
| SUT-MA-059 | CRD-0601-0434-241218-027_Voice_CPR_END |
| SUT-MA-060 | CRD-0601-0434-241218-028_ErrCodePrompt |
| SUT-MA-061 | CRD-0601-0434-241218-029_playSoftwareVersion |
| SUT-MA-062 | CRD-0601-0434-241218-030_playHardwareVersion |
| SUT-MA-063 | CRD-0601-0434-241218-031_InfoReqOpTask |
| SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
| SUT-MA-065 | CRD-0601-0434-241218-033_CPRSettingPasswordInput |
| SUT-MA-066 | CRD-0601-0434-241218-034_CommSelectTask |
| SUT-MA-067 | CRD-0601-0434-241218-035_SD_Copy_Patient_Data |
| SUT-MA-068 | CRD-0601-0434-241218-036_Init_Config_and_Operation_Data |
| SUT-MA-069 | CRD-0601-0434-241218-037_VOC_SD_to_QSPI |
| SUT-MA-070 | CRD-0601-0434-241218-038_FactoryPromptTask |
| SUT-MA-071 | CRD-0601-0434-241218-039_FactoryDefibCtrlTest |
| SUT-MA-072 | CRD-0601-0434-241218-040_FactoryAdultSelectPrompt |
| SUT-MA-073 | CRD-0601-0434-241218-041_FactoryPediatricSelectPrompt |
| SUT-MA-074 | CRD-0601-0434-241218-042_SDCardSysUpdate |
| SUT-MA-075 | CRD-0601-0434-241218-043_FM_Detail_Short_Set |
| SUT-MA-076 | CRD-0601-0434-241218-044_FM_Language_set |
| SUT-MA-077 | CRD-0601-0434-241218-045_PwrOnSelfTest |
| SUT-MA-078 | CRD-0601-0434-241218-046_POST_DisarmTest |
| SUT-MA-079 | CRD-0601-0434-241218-047_StateLCDDisplay |
| SUT-MA-080 | CRD-0601-0434-241218-048_PSTOpTask |
| SUT-MA-081 | CRD-0601-0434-241218-049_ReadTemperature |
| SUT-MA-082 | CRD-0601-0434-241218-050_ChargeMsgCtrl |
| SUT-MA-083 | CRD-0601-0434-241218-051_DefiMsgCtrl |
| SUT-MA-084 | CRD-0601-0434-241218-052_DefiEnergyCtrl |
| SUT-MA-085 | CRD-0601-0434-241218-053_batteryStatusfromVoltage |
| SUT-MA-086 | CRD-0601-0434-241218-054_ChargeDisarmTest |
| SUT-MA-087 | CRD-0601-0434-241218-055_PIC_LED_COLOR_SET |
| SUT-MA-088 | CRD-0601-0434-241218-056_AlarmBuzzerThreeTimes |
| SUT-MA-089 | CRD-0601-0434-241218-057_RedInfLEDTask |
| SUT-MA-090 | CRD-0601-0434-241218-058_ShockRampTask |
| SUT-MA-091 | CRD-0601-0434-241219-001_InfoBlueRampTask |
| SUT-MA-092 | CRD-0601-0434-241219-002_PushInfoMsgTask |
| SUT-MA-093 | CRD-0601-0434-241219-003_ShkKeyTestMsgTask |
| SUT-MA-094 | CRD-0601-0434-241219-004_HVWrCmdProc |
| SUT-MA-095 | CRD-0601-0434-241219-005_HVRdCmdProc |
| SUT-MA-096 | CRD-0601-0434-241219-006_IntADC_Init |
| SUT-MA-097 | CRD-0601-0434-241219-007_IntADCPwrOn |
| SUT-MA-098 | CRD-0601-0434-241219-008_IntADCPwrOff |
| SUT-MA-099 | CRD-0601-0434-241219-009_IntADCGet |
| SUT-MA-100 | CRD-0601-0434-241219-010_SD_Card_Update |
| SUT-MA-101 | CRD-0601-0434-241219-011_SD_Copy_Patient_Data_Load |
| SUT-MA-102 | CRD-0601-0434-241219-012_HVRdCmd |
| SUT-MA-103 | CRD-0601-0434-241219-013_HVWrCmd |
| SUT-MA-104 | CRD-0601-0434-241219-014_PwrCtrlGet |
| SUT-MA-105 | CRD-0601-0434-241219-015_PwrCtrlPut |
| SUT-MA-106 | CRD-0601-0434-241219-016_CommBurstTxMode |
| SUT-MA-107 | CRD-0601-0434-241219-017_CommHostMode |
| SUT-MA-108 | CRD-0601-0434-241219-018_ExtADC_Init |
| SUT-MA-109 | CRD-0601-0434-241219-019_ExtADC_Enable |
| SUT-MA-110 | CRD-0601-0434-241219-020_ExtADC_Get |
| SUT-MA-111 | CRD-0601-0434-241219-021_ExtRTC_Get |
| SUT-MA-112 | CRD-0601-0434-241219-022_ExtRTC_Put |
| SUT-MA-113 | CRD-0601-0434-241219-023_LED_InfoRED |
| SUT-MA-114 | CRD-0601-0434-241219-024_LED_InfoBLUE |
| SUT-MA-115 | CRD-0601-0434-241219-025_LED_ShockRamp |
| SUT-MA-116 | CRD-0601-0434-241219-026_LED_Power_Switch_Ramp |
| SUT-MA-117 | CRD-0601-0434-241219-027_LED_Patient_Contact_Ban |
| SUT-MA-118 | CRD-0601-0434-241219-028_LED_Pads_Indication_Ramp |
| SUT-MA-119 | CRD-0601-0434-241219-029_LED_Patient_Contact_Good |
| SUT-MA-120 | CRD-0601-0434-241219-030_Buzzer_Timer_Init |
| SUT-MA-121 | CRD-0601-0434-241219-031_Buzzer_ON |
| SUT-MA-122 | CRD-0601-0434-241219-032_Buzzer_OFF |
| SUT-MA-123 | CRD-0601-0434-241219-033_Audio_VolumeSet |
| SUT-MA-124 | CRD-0601-0434-241219-034_AudioPlayCtrl |
| SUT-MA-125 | CRD-0601-0434-241219-035_KEYPAD_InfKeyChkOn |
| SUT-MA-126 | CRD-0601-0434-241219-036_KEYPAD_ShkKeyChkOn |
| SUT-MA-127 | CRD-0601-0434-241219-037_KEYPAD_PwrKeyChkOn |
| SUT-MP-001 | CRD-0601-0434-241216-001_RunModeTask |
| SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
| SUT-MP-003 | CRD-0601-0434-241216-003_LoopWithSleepModeForever |
| SUT-MP-004 | CRD-0601-0434-241216-004_ChgSleepMode |
| SUT-MP-005 | CRD-0601-0434-241216-005_SleepModeTask |
| SUT-MP-006 | CRD-0601-0434-241216-006_IICIsr |
| SUT-MP-007 | CRD-0601-0434-241216-007_IICEnable |
| SUT-MP-008 | CRD-0601-0434-241216-008_IICDisable |
| SUT-MP-009 | CRD-0601-0434-241216-009_ds2431_init |
| SUT-MP-010 | CRD-0601-0434-241216-010_OW_read_byte |
| SUT-MP-011 | CRD-0601-0434-241216-011_ds2431_read_eeprom |
| SUT-MP-012 | CRD-0601-0434-241216-012_OW_write_byte |
| SUT-MP-013 | CRD-0601-0434-241216-013_ds2431_write_eeprom |
| SUT-MP-014 | CRD-0601-0434-241216-014_PadsOWReset |
| SUT-MP-015 | CRD-0601-0434-241216-015_LCDUpdateDuringSleep |
| SUT-MP-016 | CRD-0601-0434-241216-016_LED_UPDATE |
| SUT-MP-017 | CRD-0601-0434-241216-017_Set_LCD_Low_When_True |
| SUT-MP-018 | CRD-0601-0434-241216-018_Set_LCD_High_When_True |
| SUT-TP-001 | CRD-0601-0434-241213-001_Charge |
| SUT-TP-002 | CRD-0601-0434-241213-002_BiteDischgOff |
| SUT-TP-003 | CRD-0601-0434-241213-003_Discharge |
| SUT-TP-004 | CRD-0601-0434-241213-004_Disarm |
| SUT-TP-005 | CRD-0601-0434-241213-005_FirstPhaseDischgOn |
| SUT-TP-006 | CRD-0601-0434-241213-006_FirstPhaseDischgOff |
| SUT-TP-007 | CRD-0601-0434-241213-007_SecondPhaseDischgOn |
| SUT-TP-008 | CRD-0601-0434-241213-008_SecondPhaseDischgOff |
| SUT-TP-009 | CRD-0601-0434-241213-009_AdcInit |
| SUT-TP-010 | CRD-0601-0434-241213-010_AdcSetChannel |
| SUT-TP-011 | CRD-0601-0434-241213-011_AdcRead |
| SUT-TP-012 | CRD-0601-0434-241213-012_main |
| SUT-TP-013 | CRD-0601-0434-241213-013_RxISR |
| SUT-TP-014 | CRD-0601-0434-241213-014_TxISR |

Software integration and integration testing

The results of component integration and interface testing are recorded in the Software Integration Test Report

Table. Software Integration Test Report List

| SIT ID | Document Title |
| --- | --- |
| SIT-EA-01 | [CRD-0601-0434-250804-001] ECG Data Collection Test Report |
| SIT-EA-02 | [CRD-0601-0434-250804-002] ECG Data Filtering Test Report |
| SIT-EA-03 | [CRD-0601-0434-250804-003] ECG Data Analyzing Test Report |
| SIT-SE-01 | [CRD-0601-0434-250807-001] Patient Mode–Based Energy Charging Start Test Report |
| SIT-SE-02 | [CRD-0601-0434-250807-002] ECG Analysis–Based Energy Delivery Decision and User Guidance Test Report |
| SIT-SE-03-01 | [CRD-0601-0434-250807-003] Charge/Discharge Step-wise Command Transmitting Test Report |
| SIT-SE-03-02 | [CRD-0601-0434-250807-004] Therapy B/D Charging Status Monitoring Test Report |
| SIT-SE-03-03 | [CRD-0601-0434-250807-005] Therapy B/D Communication Timeout Test Report |
| SIT-PD-01 | [CRD-0601-0434-250811-001] Pads Connector Connection Status Test Report |
| SIT-PD-02 | [CRD-0601-0434-250811-002] Pads Internal Memory Verification Test Report |
| SIT-PD-03 | [CRD-0601-0434-250811-003] Pad Storage Condition Assessment Verification Test Report |
| SIT-PD-04 | [CRD-0601-0434-250811-004] Pads Attachment Status Realtime Monitoring Test Report |
| SIT-BM-01 | [CRD-0601-0569-250814-001] Battery Level Classification and State Detection Integration Test Report |
| SIT-BM-02 | [CRD-0601-0569-250814-002] Residual Battery Capacity Calculation Test Report |
| SIT-BM-03 | [CRD-0601-0569-250814-003] Low Battery Status Detection Test Report |
| SIT-BM-04 | [CRD-0601-0569-250814-004] Battery Abnormality Diagnosis Verification Test Report |
| SIT-ST-01 | [CRD-0601-0569-250818-001] POST Test Report |
| SIT-ST-02 | [CRD-0601-0569-250818-002] POST Cycle Execution Test Report |
| SIT-ST-03 | [CRD-0601-0569-250818-003] BIT Test Report |
| SIT-UI-01-01 | [CRD-0601-0569-250821-001] Main ARM UI Output Integration Test Report |
| SIT-UI-01-02 | [CRD-0601-0569-250821-002] Device Status Notification Test (Low Battery Status) |
| SIT-UI-01-03 | [CRD-0601-0569-250821-003] Device Status Notification Test (System Error Status) |
| SIT-UI-02-01 | [CRD-0601-0569-250822-001] CPR Guide Mode Entry Integration Test Report |
| SIT-UI-02-02 | [CRD-0601-0569-250822-002] Post-Defibrillation CPR Guidance in CPR Guide Mode Integration Test Report |
| SIT-UI-02-03 | [CRD-0601-0569-250822-003] CPR Guidance Transition After Internal Discharge Integration Test Report |
| SIT-UI-03 | [CRD-0601-0569-250825-001] Pad Attachment Instructions Test Report |
| SIT-UI-04 | [CRD-0601-0569-250825-002] Device Startup Guidance Test Report |
| SIT-UI-05 | [CRD-0601-0569-250825-003] Patient Mode Determination Test Report |
| SIT-UI-06 | [CRD-0601-0569-250825-004] CPR Activity Detection Test Report |
| SIT-UI-07-01 | [CRD-0601-0569-250828-001] Voice Output-based Integrated CPR Guidance Test Report |
| SIT-UI-07-02 | [CRD-0601-0569-250828-002] Operation Mode-based Integrated CPR Guidance Test Report |
| SIT-UI-08 | [CRD-0601-0569-250828-003] Pads Attachment Voice Prompt Setup Test Report |
| SIT-UI-09 | [CRD-0601-0569-250828-004] Main ARM UI Input Integration Test Report |
| SIT-DM-01 | [CRD-0601-0434-250904-001] Norflash Memory Data Storage Test Report |
| SIT-DM-02 | [CRD-0601-0434-250904-002] ECG Data Storage Test Report |
| SIT-DM-03 | [CRD-0601-0434-250904-003] Treatment Record Extraction Test Report |
| SIT-DV-01 | [CRD-0601-0569-250911-001] Wake-up Status Check and Initialization Test Report |
| SIT-DV-02 | [CRD-0601-0569-250911-002] Real-time Power-off Detection Monitoring Test Report |
| SIT-DV-03-01 | [CRD-0601-0569-250911-003] System Error Detection and Error Type Classification Test Report |
| SIT-DV-03-02 | [CRD-0601-0569-250911-004] Error Code Format and Range Validation with Retry Handling Test Report |
| SIT-DV-03-03 | [CRD-0601-0569-250911-005] User Notification and Power Shutdown Control for Persistent Error Conditions Test Report |
| SIT-DV-04 | [CRD-0601-0434-250918-001] Environment Noise Analysis and Voice Prompt Volume Control Test Report |
| SIT-DV-05 | [CRD-0601-0434-250918-002] CPR Guidance for System Error Conditions Test Report |
| SIT-DV-06 | [CRD-0601-0434-250918-003] System Mode Determination Test Report |
| SIT-DV-07 | [CRD-0601-0434-250925-001] Administrator Mode Entry Test Report |
| SIT-DV-08 | [CRD-0601-0434-250925-002] CPR Guidance Setup Update Test Report |
| SIT-DV-09-01 | [CRD-0601-0434-250925-003] Factory Mode Device Setup Test Report |
| SIT-DV-09-02 | [CRD-0601-0434-250925-004] Factory Mode Device Initialization Setup Test Report |
| SIT-DV-10 | [CRD-0601-0569-251002-001] Factory Mode SD Card-Based Device Setup Test Report |
| SIT-DV-11 | [CRD-0601-0569-251002-002] IrDA-Based Device Setup Change Test Report |
| SIT-DV-12 | [CRD-0601-0569-251002-003] Post-Power-on Idle State Monitoring Test Report |
| SIT-CA-01 | [CRD-0601-0434-251009-001] SD Card Read/Write Operations Test Report |
| SIT-CA-02 | [CRD-0601-0434-251009-002] Therapy PIC-High Voltage Module Communication Test Report |
| SIT-CA-03 | [CRD-0601-0434-251009-003] Main PIC–EEPROM Control Module Communication Test Report |
| SIT-CA-04 | [CRD-0601-0434-251009-004] IrDA Communication Test Report |
| SIT-CM-01 | [CRD-0601-0434-251010-001] Main PIC I2C Communication Test Report |
| SIT-UM-01 | [CRD-0601-0569-251016-001] Main PIC UI Input Integration Test Report |
| SIT-UM-02-01 | [CRD-0601-0569-251016-002] Main PIC UI Output Control (LED) Test Report |
| SIT-UM-02-02 | [CRD-0601-0569-251016-003] Main PIC UI Output Control (LCD) Test Report |
| SIT-EE-01 | [CRD-0601-0569-251023-001] Pads and Battery Information Test Report |
| SIT-PM-01 | [CRD-0601-0569-251023-002] Main ARM Power Activation Test Report |
| SIT-PM-02 | [CRD-0601-0569-251023-003] Main ARM Status Monitoring Test Report |
| SIT-PM-03 | [CRD-0601-0569-251023-004] Main ARM Power Deactivation Test Report |
| SIT-PM-04 | [CRD-0601-0569-251030-001] Low Power Standby Mode Transition Test Report |
| SIT-PM-05 | [CRD-0601-0569-251030-002] Main PIC RTC Control Test Report |
| SIT-CT-01 | [CRD-0601-0569-251030-003] Main ARM–Therapy PIC Integration Communication Test Report |
| SIT-HV-01 | [CRD-0601-0434-251105-001] High Voltage Module Energy Charging Control and Status Check Test Report |
| SIT-HV-02 | [CRD-0601-0434-251105-002] External Shock Energy Discharge Test Report |
| SIT-HV-03 | [CRD-0601-0434-251105-003] Internal Shock Energy Discharge Test Report |
| SIT-HV-04 | [CRD-0601-0434-251105-004] External Defibrillation Biphasic Waveform Integration Test Report |

Software System Testing

The results of testing for all requirements and the verification of risk control measures are recorded in the Software System Test Report

Table. Software System Test Report List

| SST ID | Title |
| --- | --- |
| SST-01 | [CRD-0601-0434-260105-001] ECG Measurement and Analysis System Test Report |
| SST-02 | [CRD-0601-0434-260108-001] Shock Energy Delivery (External Discharge) System Test Report |
| SST-03-01 | [CRD-0601-0569-260112-001] CPR Guidance-1 System Test Report |
| SST-03-02 | [CRD-0601-0569-260112-002] CPR Guidance-2 System Test Report |
| SST-03-03 | [CRD-0601-0569-260112-003] CPR Guidance-3 (Pads Attachment Guide - detailed mode) System Test Report |
| SST-03-04 | [CRD-0601-0569-260112-004] CPR Guidance-3 (Pads Attachment Guide - shortened mode) System Test Report |
| SST-03-05 | [CRD-0601-0569-260112-005] AHA Guidelines Compliance System Test Report |
| SST-04 | [CRD-0601-0434-260115-001] Device Status Monitoring System Test Report |
| SST-05-01 | [CRD-0601-0569-260116-001] Pads Management (Defibrillator Pads Connector) System Test Report |
| SST-05-02 | [CRD-0601-0569-260116-002] Pads Management (Status and Attachment) System Test Report |
| SST-06 | [CRD-0601-0434-260120-001] Device Operation Status Guidance System Test Report |
| SST-07 | [CRD-0601-0569-260122-001] Shock Energy Delivery (Internal Discharge) System Test Report |
| SST-08-01 | [CRD-0601-0569-260123-001] Real-time Monitoring System for High Voltage Capacitor Charging Voltage Test Report |
| SST-08-02 | [CRD-0601-0569-260123-002] High Voltage Capacitor Overcharge Monitoring System Test Report |
| SST-09 | [CRD-0601-0434-260127-001] POST(Power-On Self-Test) System Test |
| SST-10-01 | [CRD-0601-0569-260129-001] PST(Periodic Self-Test) System Test |
| SST-10-02 | [CRD-0601-0569-260129-002] PST(Periodic Self-Test) Error Detection System Test Report |
| SST-11 | [CRD-0601-0434-260202-001] BIT(Battert Insertion Test) System Test |
| SST-12 | [CRD-0601-0434-260202-002] Patient Treatment Data Storage and Export System Test Report |
| SST-13 | [CRD-0601-0434-260202-003] Device Unavailability Guidance System Test Report |
| SST-14 | [CRD-0601-0569-260209-001] Error CPR Guidance System Test |
| SST-15 | [CRD-0601-0569-260210-001] CPR Guidance Settings System Test |
| SST-16-01 | [CRD-0601-0569-260211-001] Factory Mode Settings System Test |
| SST-16-02 | [CRD-0601-0569-260211-002] Factory Mode Settings Change System Test |
| SST-16-03 | [CRD-0601-0569-260211-003] Factory Mode Settings Change via IrDA and SD Card System Test Report |
| SST-17 | [CRD-0601-0434-260213-001] AED Standby Mode Maintenance System Test Report |

Software Release for utilization at a SYSTEM level

The software release was performed according to the Software Release Guidelines [CUI-731.002].

Software released Versions

A Software release is the distribution of the final version or the newest version of a software application. A software release may be public or private and generally signifies the unveiling of a new or upgraded version of the application.

To be successful, software release management requires careful planning and execution, with customer-centric goals, time and budget efficiency, and minimal negative impact on customers.

Refer to the table below for information on SP1 Series Software releases.

| No. | Version (Release date) | Reason of Change | Classify of change -Improvement -Adaptive -Correction |
| --- | --- | --- | --- |
| 1 | Ver1.00 (2011.4.25) | First Release | First Release |
| 2 | Ver 1.10 (2012.4.2) | Battery level detection algorithm modified. | Correction & Improvement |
| 3 | Ver 1.20 (2013.9.6) | Added following voice prompts for Japanese Upgrading Analyze Module | Improvement |
| 4 | Ver 1.30 (2014.6.3) | Changing Shock button delay Time (Applied only for Japanese exports.) | Improvement |
| 5 | Ver 1.40 (2015.4.14) | Improved System Monitor Module. | Improvement |
| 6 | Ver 1.42 (2018.11.06) | Add an additional battery status display stage: define alarm level | Improvement |
| 7 | Ver 1.43 (2020.02.10) | Improved Pads status check function | Improvement |
| 8 | Ver 1.50 (2021.03.15) | Rectifying software for the error fix (UK Only) | Improvement |
| 9 | Ver 1.51 (2022.01.14) | Improper voice prompt delivery for pads expiry date when turned on. No configuration change was made when updated from below version 1.42. | Improvement |
|  | Ver 1.52 (2022.01.14) | Improper voice prompt delivery for pads expiry date when turned on. | Improvement |
| 10 | Ver 2.00 (2020.07.14) | First Release - Changed Main Microprocessor | Improvement |
| 11 | Ver 2.02 (2022.01.07) | Update for fixing bugs | Improvement |
| 12 | Ver 2.03 (2022.09.06) | Changed language packs to improve sound quality | Improvement |

Software Risk Management PROCESS

The risk management process consists of the following four steps.

- risk analysis to identify potential hazards and causes contributing to a hazardous situation;

- risk assessment to evaluate the possible consequences of a hazard;

- Risk controls applied to mitigate risk; and

- After risk mitigation, residual risk assessment to determine whether the risk is acceptable

Risk Analysis

Identifies software items that could contribute the hazardous situations identified in the medical device risk analysis activity of ISO 14971.

The potential causes of the software items identified above contributing to the hazardous situation shall be identified. Potential causes to consider include:

- Incorrect or incomplete specification of functionality;

- Software defects in the identified SOFTWARE ITEM functionality;

- Hardware failures or other software defects that could result in unpredictable software operation; and

- Reasonably foreseeable misuse.

Refer to Cl 5.1.7.1 of Software Development Plan [CRD-0601-PP-SDP-01].

Risk Evaluation

Risk evaluation is an iterative process. Risks are assessed before risk control measures are applied and re-evaluated each time risk control measures are applied.

The severity of harm is evaluated according to Table 9.2-1 below.

| Five quantitative severity levels |  |  |
| --- | --- | --- |
| Common terms | Rating | Possible description |
| Catastrophic | 5 | Results in patient death |
| Critical | 4 | Results in permanent impairment or life threatening injury |
| Moderate | 3 | Results in injury or impairment requiring professional medical intervention |
| Minor | 2 | Results in temporary injury or impairment not requiring professional medical intervention |
| Negligible | 1 | Inconvenience or temporary discomfort |

Table 9.2-1 Severity of Harm

The probability of a hazard is evaluated according to Table 9.2-2 below.

| Five quantitative severity levels |  |  |
| --- | --- | --- |
| Common terms | Rating | Possible description |
| Catastrophic | 5 | Results in patient death |
| Critical | 4 | Results in permanent impairment or life-threatening injury |
| Moderate | 3 | Results in injury or impairment requiring professional medical intervention |
| Minor | 2 | Results in temporary injury or impairment not requiring professional medical intervention |
| Negligible | 1 | Inconvenience or temporary discomfort |

Table 9.2-2 Probability of Harm

Risk Control

If a software item can contribute to a hazardous situation, risk control measures are defined and documented in accordance with ISO 14971 for each case documented in the risk management file. Control methods for software risk management include risk control using hardware, risk control using software, risk control using work environment, and risk control using Instructions for Use.

The implementation of each risk control measure should be verified, it is determined whether it could create a new hazardous situation, and the risk control measures are reviewed.

Evaluation of overall residual risk

The following table calculates the software safety level by multiplying the probability and severity according to the risks arising from the software system.

| Risk evaluation matrix for CU-SP1, CU-SP1 PLUS |  |  | severity levels |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Negligible | Minor | Moderate | Critical | Catastrophic |
| (Probability × Severity) |  | Rating | 1 | 2 | 3 | 4 | 5 |
| Probability levels | Frequent | 5 | 5 | 10 | 15 | 20 | 25 |
|  | Probable | 4 | 4 | 8 | 12 | 16 | 20 |
|  | Occasional | 3 | 3 | 6 | 9 | 12 | 15 |
|  | Remote | 2 | 2 | 4 | 6 | 8 | 10 |
|  | Improbable | 1 | 1 | 2 | 3 | 4 | 5 |

| Risk Level | Description |
| --- | --- |
| Acceptable Risk (Yellow zone + Green zone) | Hazards whose initial risk rating falls Risk (Yellow zone + Green zone) are deemed acceptable risks for the users, patients and environment of the CU-SP1 & CU-SP1 PLUS given the benefits obtained from the use of the device. Where practical, risk control measures for this type of hazard will be included in the design of the device to further reduce the risk of harm. All risks have to be reduced as far as possible, regardless of any acceptability. |
| Unacceptable Risk (Red zone) | Hazards whose initial risk rating exceeds Risk (Red zone) must apply sufficient risk control measures to reduce the risk associated with that hazard to an acceptable level. Hazards whose residual risk rating still exceeds the threshold risk must apply an additional risk control measure which reduces the residual risk rating |

Changes to the medical device software are analyzed to determine whether additional potential causes appear to contribute to a hazardous situation or if additional software risk control measures are required. Change analysis determines whether a software modification could disrupt existing risk control measures. Based on this analysis, risk management activities are performed.

FMEA Table

For software-related risk analysis, refer to the FMEA table provided in the FMEA Report (CRD-0601-DI-RMF-01).

Software Configuration Management

All identified problems are handled according to the procedures defined in the Software Development Guidelines [CUI-731.001].

Software Problem Resolution

All identified problems are handled according to the procedures defined in the Software Maintenance Guidelines [CUI-731.004].

- Appendix 1. IEC 62304 Checklist

| Clause | Requirement | Result-Remark | Verdict |
| --- | --- | --- | --- |
| 4 | GENERAL REQUIREMENTS |  | — |
| 4.1 | [A, B, C] Quality management system |  | — |
|  | The MANUFACTURER of MEDICAL DEVICE SOFTWARE demonstrates the ability to provide MEDICAL DEVICE SOFTWARE that consistently meets customer requirements and applicable regulatory requirements | ISO 13485 Certificate (9807-2017-AQ-KOR-NA-PS) | P |
| 4.2 | [A, B, C] RISK MANAGEMENT |  | — |
|  | The MANUFACTURER applies a RISK MANAGEMENT PROCESS complying with ISO 14971 | Refer to Risk Management Procedure (CUP-410) | P |
| 4.3 | [A, B, C] Software safety classification |  | — |
|  | - The MANUFACTURER assigns to each SOFTWARE SYSTEM a software safety class according to the RISK of HARM to the patient, operator, or other people resulting from a HAZARDOUS SITUATION to which the SOFTWARE SYSTEM can contribute in a worst-case-scenario | Software safety class C is applied. Refer to Software Design Specification (CRD-0601-DI-SDS-01) | P |
|  | The SOFTWARE SYSTEM is software safety class A if: |  | N/A |
|  | - the SOFTWARE SYSTEM does not contribute to a HAZARDOUS SITUATION; or - the SOFTWARE SYSTEM contributes to a HAZARDOUS SITUATION which does not result in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM | Software safety class A is not applicable | N/A |
|  | The SOFTWARE SYSTEM is software safety class B if: |  | N/A |
|  | - the SOFTWARE SYSTEM contributes to a HAZARDOUS SITUATION which results in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is non-SERIOUS INJURY | Software safety class B is not applicable | N/A |
|  | The SOFTWARE SYSTEM is software safety class C if: |  | — |
|  | - the SOFTWARE SYSTEM contributes to a HAZARDOUS SITUATION which results in unacceptable RISK after consideration of RISK CONTROL measures external to the SOFTWARE SYSTEM and the resulting possible HARM is death or SERIOUS INJURY | Software safety class C is applied | P |
|  | For a SOFTWARE SYSTEM initially classified as software safety class B or C, the MANUFACTURER has implemented additional RISK CONTROL measures external to the SOFTWARE SYSTEM and subsequently has assigned a new software safety classification to the SOFTWARE SYSTEM | New software safety classification is not assigned. | P |
|  | - The MANUFACTURER documents the software safety class assigned to each SOFTWARE SYSTEM in the RISK MANAGEMENT FILE | Class C described in Risk Management Report (CRD-0601-VV-RMR-01) and Software Design Specification (CRD-0601-DI-SDS-01) | P |
|  | - When a SOFTWARE SYSTEM is decomposed into SOFTWARE ITEMS, and when a SOFTWARE ITEM is decomposed into further SOFTWARE ITEMS, such SOFTWARE ITEMS inherit the software safety classification of the original SOFTWARE ITEM (or SOFTWARE SYSTEM) unless the MANUFACTURER documents a rationale for classification into a different software safety class | The classification of SOFTWARE ITEMS is identical to that of a SOFTWARE SYSTEM. Refer to Cl 5 of Software Design Specification (CRD-0601-DI-SDS-01) | P |
|  | A rationale explains how the new SOFTWARE ITEMS are segregated so that they may be classified separately | See above. | N/A |
|  | - The MANUFACTURER documents the software safety class of each SOFTWARE ITEM if that class is different from the class of the SOFTWARE ITEM from which it was created by decomposition | The classification of SOFTWARE ITEMS is identical to that of a SOFTWARE SYSTEM. | N/A |
|  | - When applied to a group of SOFTWARE ITEMS, the MANUFACTURER uses the PROCESSES and TASKS which are required by the classification of the highest-classified SOFTWARE ITEM in the group unless the MANUFACTURER documents in the RISK MANAGEMENT FILE a rationale for using a lower classification | The classification of SOFTWARE ITEMS is identical to that of a SOFTWARE SYSTEM. | N/A |
|  | - Class C requirements apply for each SOFTWARE SYSTEM, until a software safety class is assigned | Software safety classification is assigned to Class C | N/A |
| 4.4 | [A, B, C] LEGACY SOFTWARE |  | - |
|  | Clauses 5 through 9 have applied to demonstrate the compliance of LEGACY SOFTWARE | No legacy software | N/A |
|  | As alternative, clauses 4.4.2 through 4.4.5 have applied to demonstrate the compliance of LEGACY SOFTWARE | See above | N/A |
| 4.4.2 | [A, B, C] RISK MANAGEMENT ACTIVITIES |  | N/A |
|  | The MANUFACTURER: |  | N/A |
|  | - assesses any feedback, including post-production information, on LEGACY SOFTWARE regarding incidents and / or near incidents, both from inside its own organization and / or from users | No legacy software | N/A |
|  | - performs RISK MANAGEMENT ACTIVITIES associated with continued use of the LEGACY SOFTWARE | See above | N/A |
|  | Considering the following aspects: |  | — |
|  | - integration of the LEGACY SOFTWARE in the overall MEDICAL DEVICE architecture | See above | N/A |
|  | - continuing validity of RISK CONTROL measures, implemented as part of the LEGACY SOFTWARE | See above | N/A |
|  | - identification of HAZARDOUS SITUATIONS associated with the continued use of the LEGACY SOFTWARE | See above | N/A |
|  | - identification of potential causes of the LEGACY SOFTWARE contributing to a HAZARDOUS SITUATIONS | See above | N/A |
|  | - definition of RISK CONTROL measures for each potential cause of the LEGACY SOFTWARE contributing to a HAZARDOUS SITUATIONS | See above | N/A |
| 4.4.3 | [A, B, C] Gap analysis |  | N/A |
|  | Based on the software safety class of the LEGACY SOFTWARE, the MANUFACTURER performs a gap analysis of available DELIVERABLES against those required according to 5.2, 5.3, 5.7, and Clause 7 | No legacy software | N/A |
|  | - The MANUFACTURER assesses the continuing validity of available DELIVERABLES | See above | N/A |
|  | - Where gaps are identified, the MANUFACTURER EVALUATES the potential reduction in RISK resulting from the generation of the missing DELIVERABLES and associated ACTIVITIES | See above | N/A |
|  | - Based on this evaluation, the MANUFACTURER determines the DELIVERABLES to be created and associated ACTIVITIES to be performed | See above | N/A |
|  | SOFTWARE SYSTEM test records are the minimum DELIVERABLES to be created | See above | N/A |
| 4.4.4 | [A, B, C] Gap closure activities |  | N/A |
|  | - The MANUFACTURER establishes and executes a plan to generate the identified DELIVERABLES | No legacy software | N/A |
|  | Objective evidence have used to generate required DELIVERABLES without performing ACTIVITIES required by 5.2, 5.3, 5.7 and Clause 7 | See above | N/A |
|  | - The plan addresses the use of the problem resolution PROCESS for handling problems detected in the LEGACY SOFTWARE and DELIVERABLES in accordance with Clause 9 | See above | N/A |
|  | - Changes to the LEGACY SOFTWARE have performed in accordance with Clause 6. | See above | N/A |
| 4.4.5 | [A, B, C] Rationale for use of LEGACY SOFTWARE |  | — |
|  | The MANUFACTURER documents the VERSION of the LEGACY SOFTWARE together with a rationale for the continued use of the LEGACY SOFTWARE | No legacy software | N/A |
| 5 | SOFTWARE DEVELOPMENT PROCESS |  | — |
| 5.1 | Software development planning |  | — |
| 5.1.1 | [A, B, C] The MANUFACTURER establishes a software development plan (or plans) for conducting the ACTIVITIES of the software development PROCESS appropriate to the scope, magnitude, and software safety classifications of the SOFTWARE SYSTEM to be developed. | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The SOFTWARE DEVELOPMENT LIFE CYCLE MODEL is either fully- defined or be referenced in the plan (or plans). | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The plan addresses the following: |  | P |
|  | - the PROCESSES to be used in the development of the SOFTWARE SYSTEM | See above | P |
|  | - the DELIVERABLES (includes documentation) of the ACTIVITIES and TASKS | See above | P |
|  | - TRACEABILITY between SYSTEM requirements, software requirements, SOFTWARE SYSTEM test, and RISK CONTROL measures implemented in software | See above | P |
|  | - software configuration and change management, including SOUP CONFIGURATION ITEMS and software used to support development | See above | P |
|  | - software problem resolution for handling problems detected in the MEDICAL DEVICE SOFTWARE, DELIVERABLES and ACTIVITIES at each stage of the life cycle | See above | P |
| 5.1.2 | [A, B, C] The MANUFACTURER updates the plan, as appropriate, as development proceeds | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.3 | [A, B, C] Software development plan reference to SYSTEM design and development |  | P |
|  | - As inputs for software development, SYSTEM requirements are referenced in the software development plan by the MANUFACTURER | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - In the software development plan, the MANUFACTURER includes or references procedures for coordinating the software development with the system development necessary to satisfy 4.1 (such as system integration, verification, and validation) | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.4 | [C] Associated with the development of SOFTWARE ITEMS of class C, in the software development plan are included or referenced: |  | P |
|  | - standards | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - methods | See above | P |
|  | - tools | See above | P |
| 5.1.5 | [B, C] The MANUFACTURER includes or references in the software development plan, a plan to integrate the SOFTWARE ITEMS (including SOUP) and performs testing during integration | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.6 | [A, B, C] In the software development plan, the following VERIFICATION information are included or referenced: |  | P |
|  | - DELIVERABLES requiring VERIFICATION | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - the required VERIFICATION TASKS for each life cycle ACTIVITY | See above. | P |
|  | - milestones at which the DELIVERABLES are VERIFIED | See above. | P |
|  | - the acceptance criteria for VERIFICATION of the DELIVERABLES | See above. | P |
| 5.1.7 | [A, B, C] In the software development plan the MANUFACTURER includes or references a plan to conduct the ACTIVITIES and TASKS of the software RISK MANAGEMENT PROCESS, including the management of RISKS relating to SOUP | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.8 | [A, B, C] In the software development plan the MANUFACTURER includes or references information about the documents to be produced during the software development life cycle | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | For each identified document or type of document the following information has included or referenced: |  | P |
|  | - title, name or naming convention | See above | P |
|  | - purpose | See above | P |
|  | - procedures and responsibilities for development, review, approval and modification | Refer to Product Development Work Procedure (CUP-731) and Software Procedure (CUI-731.001) | P |
| 5.1.9 | [A, B, C] The MANUFACTURER includes or references software configuration management information in the software development plan | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The software configuration management information includes or references: |  | P |
|  | - the classes, types, categories or lists of items to be controlled | See above | P |
|  | - the software configuration management ACTIVITIES and TASKS | See above | P |
|  | - the organization(s) responsible for performing software configuration management and ACTIVITIES | See above | P |
|  | - their relationship with other organizations, such as software development or maintenance | See above | P |
|  | - when the items are to be placed under configuration control | See above | P |
|  | - when the problem resolution PROCESS is to be used | See above | P |
| 5.1.10 | [B, C] The items to be controlled include tools, items or settings, used to develop the MEDICAL DEVICE SOFTWARE, which could impact the MEDICAL DEVICE SOFTWARE | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.11 | [B, C] The MANUFACTURER plans to place CONFIGURATION ITEMS under documented configuration management control before they are VERIFIED | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.1.12 | [B, C] In the software development plan the MANUFACTURER includes or references a procedure for: |  | P |
|  | - identifying categories of defects that may be introduced based on the selected programming technology that are relevant to their SOFTWARE SYSTEM | Refer to Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - documenting evidence that demonstrates that these defects do not contribute to unacceptable RISK | See above | P |
| 5.2 | Software requirements analysis |  | — |
| 5.2.1 | [A, B, C] For each SOFTWARE SYSTEM of the MEDICAL DEVICE, the MANUFACTURER defines and documents SOFTWARE SYSTEM requirements from the SYSTEM level requirements | Refer to Software Requirement Specification (CRD-0601-DI-SRS-01) | P |
| 5.2.2 | [A, B, C] As appropriate to the MEDICAL DEVICE SOFTWARE, the MANUFACTURER includes in the software requirements: |  | P |
|  | - functional and capability requirements | Software Requirement Specification (CRD-0601-DI-SRS-01) | P |
|  | - SOFTWARE SYSTEM inputs and outputs | See above. | P |
|  | - interfaces between the SOFTWARE SYSTEM and other SYSTEMS | See above. | P |
|  | - software-driven alarms, warnings, and operator messages | See above. | P |
|  | - SECURITY requirements | See above. | P |
|  | - user interface requirements implemented by software | See above. | P |
|  | - data definition and database requirements | See above. | P |
|  | - installation and acceptance requirements of the delivered MEDICAL DEVICE SOFTWARE at the operation and maintenance site or sites | No installation | N/A |
|  | - requirements related to methods of operation and maintenance | Software Requirement Specification (CRD-0601-DI-SRS-01) | P |
|  | - requirements related to IT-network aspects | Not use IT-Network | N/A |
|  | - user maintenance requirements | Software Requirement Specification (CRD-0601-DI-SRS-01) | P |
|  | - regulatory requirements | See above. | P |
| 5.2.3 | [B, C] The MANUFACTURER includes RISK CONTROL measures implemented in software in the requirements as appropriate to the MEDICAL DEVICE SOFTWARE | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) and Software FMEA Report (CRD-0604-DI-RMF-01) | P |
| 5.2.4 | [A, B, C] The MANUFACTURER re-EVALUATES the MEDICAL DEVICE RISK ANALYSIS when software requirements are established and update it as appropriate | The re-evaluates was not conducted because the requirements were not updated. | N/A |
| 5.2.5 | [A, B, C] The MANUFACTURER ensures that existing requirements, including SYSTEM requirements, are re- EVALUATED and updated as appropriate as a result of the software requirements analysis ACTIVITY | The re-evaluates was not conducted because the requirements were not updated. | N/A |
| 5.2.6 | [A, B, C] The MANUFACTURER verifies and documents that the software requirements: |  | P |
|  | - implement SYSTEM requirements including those relating to RISK CONTROL | Refer to Software Verification Test Report (CRD-0601-VV-SVR-01) | P |
|  | - do not contradict one another | See above. | P |
|  | - are expressed in terms that avoid ambiguity | See above. | P |
|  | - are stated in terms that permit establishment of test criteria and performance of tests | See above. | P |
|  | - can be uniquely identified | See above. | P |
|  | - are traceable to SYSTEM requirements or other source | See above. | P |
| 5.3 | Software ARCHITECTURAL design |  | — |
| 5.3.1 | [B, C] The MANUFACTURER transforms the requirements for the MEDICAL DEVICE SOFTWARE into a documented ARCHITECTURE that describes the software’s structure and identifies the SOFTWARE ITEMS | Refer to Software Design Specification (CRD-0601-DI-SDS-01) | P |
| 5.3.2 | [B, C] The MANUFACTURER develops and documents an ARCHITECTURE for the interfaces between the SOFTWARE ITEMS and the components external to the SOFTWARE ITEMS (both software and hardware), and between the SOFTWARE ITEMS | Refer to Software Design Specification (CRD-0601-DI-SDS-01) | P |
| 5.3.3 | [B, C] If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER specifies functional and performance requirements for the SOUP item that are necessary for its intended use | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Instruction (CUI-731.001) | P |
| 5.3.4 | [B, C] If a SOFTWARE ITEM is identified as SOUP, the MANUFACTURER specifies the SYSTEM hardware and software necessary to support the proper operation of the SOUP item | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Instruction (CUI-731.001) | P |
| 5.3.5 | [C] The MANUFACTURER identifies any segregation between SOFTWARE ITEMS that is necessary for RISK CONTROL, and states how to ensure that such segregation is effective | Refer to Cl 5.3.3 of Software Development File (CRD-0601-VV-SDF-01) and Cl 5.3 of Software Design Specification [CRD-0601-DI-SDS-01] | P |
| 5.3.6 | [B, C] The MANUFACTURER verifies and documents that: |  | P |
|  | - the ARCHITECTURE of the software implements SYSTEM and software requirements including those relating to RISK CONTROL | Refer to Software Verification Test Report (CRD-0601-VV-SVR-01) | P |
|  | - the software ARCHITECTURE is able to support interfaces between SOFTWARE ITEMS and between SOFTWARE ITEMS and hardware | See above. | P |
|  | - the MEDICAL DEVICE ARCHITECTURE supports proper operation of any SOUP items | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Instruction (CUI-731.001) | P |
| 5.4 | Software detailed design |  | — |
| 5.4.1 | [B, C] The MANUFACTURER subdivides the software until it is represented by SOFTWARE UNITS | Refer to Cl 8 of Software Design Specification [CRD-0601-DI-SDS-01] | P |
| 5.4.2 | [C] The MANUFACTURER documents a design with enough detail to allow correct implementation of each SOFTWARE UNIT | Refer to Cl 8 of Software Design Specification [CRD-0601-DI-SDS-01] | P |
| 5.4.3 | [C] The MANUFACTURER documents a design for any interfaces between the SOFTWARE UNIT and external components (hardware or software), as well as any interfaces between SOFTWARE UNITS, detailed enough to implement each SOFTWARE UNIT and its interfaces correctly | Refer to Cl 8 of Software Design Specification [CRD-0601-DI-SDS-01] | P |
| 5.4.4 | [C] The MANUFACTURER verifies and documents that the software detailed design: |  | P |
|  | - implements the software ARCHITECTURE | Refer to Cl 8 of Software Design Specification [CRD-0601-DI-SDS-01] | P |
|  | - is free from contradiction with the software ARCHITECTURE. | See above. | P |
| 5.5 | SOFTWARE UNIT implementation |  | — |
| 5.5.1 | [A, B, C] The MANUFACTURER implements each SOFTWARE UNIT | Refer to Cl 5 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.5.2 | [B, C] The MANUFACTURER establishes strategies, methods and procedures for verifying the SOFTWARE UNITS | Refer to Cl 5 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | Where VERIFICATION is done by testing, the test procedures are EVALUATED for adequacy | See above. | P |
| 5.5.3 | [B, C] The MANUFACTURER establishes acceptance criteria for SOFTWARE UNITS prior to integration into larger SOFTWARE ITEMS as appropriate, and ensures that SOFTWARE UNITS meet acceptance criteria | Refer to Cl 5 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.5.4 | [C] When present in the design, the MANUFACTURER includes additional acceptance criteria as appropriate for: |  | P |
|  | - proper event sequence | Refer to Cl 5 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - data and control flow | See above. | P |
|  | - planned resource allocation | See above. | P |
|  | - fault handling (error definition, isolation, and recovery) | See above. | P |
|  | - initialization of variables | See above. | P |
|  | - self-diagnostics | See above. | P |
|  | - memory management and memory overflows | See above. | P |
|  | - boundary conditions | See above. | P |
| 5.5.5 | [B, C] The MANUFACTURER performs the SOFTWARE UNIT VERIFICATION and documents the results | Refer to Software Verification Test Plan [CRD-0601-VV-SVP-01] | P |
| 5.6 | Software integration and integration testing |  | — |
| 5.6.1 | [B, C] The MANUFACTURER integrates the SOFTWARE UNITS in accordance with the integration plan | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.2 | [B, C] The MANUFACTURER verifies that the SOFTWARE UNITS have been integrated into SOFTWARE ITEMS and/or the SOFTWARE SYSTEM in accordance with the integration plan and retains records of the evidence of such verification | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.3 | [B, C] The MANUFACTURER tests the integrated SOFTWARE ITEMS in accordance with the integration plan and documents the results | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.4 | [B,C] For software integration testing, the MANUFACTURER addresses whether the integrated SOFTWARE ITEM performs as intended | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.5 | [B, C] The MANUFACTURER EVALUATES the integration test procedures for adequacy | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.6 | [B, C] When software items are integrated, the MANUFACTURER conducts REGRESSION TESTING appropriate to demonstrate that defects have not been introduced into previously integrated software | Refer to Cl 6 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.6.7 | [B, C] The MANUFACTURER: |  | P |
|  | - documents the test result (pass/fail and a list of ANOMALIES) | Refer to Software Verification Test Report (CRD-0601-VV-SVR-01) | P |
|  | - retains sufficient records to permit the test to be repeated | See above. | P |
|  | - identifies the tester | See above. | P |
| 5.6.8 | [B, C] The MANUFACTURER enters ANOMALIES found during software integration and integration testing into a software problem resolution PROCESS | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 5.7 | SOFTWARE SYSTEM testing |  | — |
| 5.7.1 | [A, B, C] Establish tests for software requirements |  | P |
|  | - The MANUFACTURER establishes and performs a set of tests, expressed as input stimuli, expected outcomes, pass/fail criteria and procedures, for conducting SOFTWARE SYSTEM testing, such that all software requirements are covered | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - The MANUFACTURER EVALUATES the adequacy of VERIFICATION strategies and test procedures. | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.7.2 | [A, B, C] The MANUFACTURER enters ANOMALIES found during software system testing into a software problem resolution PROCESS | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 5.7.3 | [A, B, C] When changes are made during SOFTWARE SYSTEM testing, the MANUFACTURER: |  | P |
|  | - repeats tests, performs modified tests or performs additional tests, as appropriate, to verify the effectiveness of the change in correcting the problem | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
|  | - conducts testing appropriate to demonstrate that unintended side effects have not been introduced | See above. | P |
|  | - performs relevant RISK MANAGEMENT ACTIVITIES as defined in 7.4 | See above. | P |
| 5.7.4 | [A, B, C] Evaluate SOFTWARE SYSTEM testing |  | P |
|  | The MANUFACTURER EVALUATES the appropriateness of VERIFICATION strategies and test procedures | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | The MANUFACTURER verifies that: |  | P |
|  | - all software requirements have been tested or otherwise VERIFIED | See above. | P |
|  | - the TRACEABILITY between software requirements and tests or other VERIFICATION is recorded | See above. | P |
|  | - test results meet the required pass/fail criteria | See above. | P |
| 5.7.5. | [A, B, C] In order to support the repeatability of tests, the MANUFACTURER documents: |  | P |
|  | - a reference to test case procedures showing required actions and expected results | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - the test result (pass/fail and a list of ANOMALIES) | See above. | P |
|  | - the version of software tested | See above. | P |
|  | - relevant hardware and software test configurations | See above. | P |
|  | - relevant test tools | See above. | P |
|  | - date tested | See above. | P |
|  | - the identity of the person responsible for executing the test and recording the test results | See above. | P |
| 5.8 | Software RELEASE for utilization at a SYSTEM level |  | — |
| 5.8.1 | [A, B, C] The MANUFACTURER ensures that all software VERIFICATION ACTIVITIES has been completed and the results EVALUATED before the software is released | Refer to Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.8.2 | [A, B, C] The MANUFACTURER documents all known residual ANOMALIES | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 5.8.3 | [B, C] The MANUFACTURER ensured that all known residual ANOMALIES have been EVALUATED to ensure that they do not contribute to an unacceptable RISK | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 5.8.4 | [A, B, C] The MANUFACTURER documented the VERSION of the MEDICAL DEVICE SOFTWARE that is being released | Refer to Software Development File (CRD-0601-VV-SDF-01) | P |
| 5.8.5 | [B, C] The MANUFACTURER documents the procedure and environment used to create the released software | Refer to Software Release Guidelines (CUI-731.003) | P |
| 5.8.6 | [B, C] The MANUFACTURER ensures that all software development plan (or maintenance plan) ACTIVITIES and TASKS are complete along with the associated documentation | All activities and tasks are checked by Phase review Refer to Software Development Guidelines (CUI-731.001) and Software Maintenance Guidelines (CUI-731.004) | P |
| 5.8.7 | [A, B, C] For at least a period of time determined as the longer of: the life time of the MEDICAL DEVICE SOFTWARE as defined by the MANUFACTURER or a time specified by relevant regulatory requirements, the MANUFACTURER archives: |  | P |
|  | - the MEDICAL DEVICE SOFTWARE and CONFIGURATION ITEMS | Refer to Cl 5.1.9 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
|  | - the documentation | See above. | P |
| 5.8.8 | [A, B, C] The MANUFACTURER establishes procedures to ensure that the released MEDICAL DEVICE SOFTWARE can be reliably delivered to the point of use without corruption or unauthorised change | Medical device software is embedded into medical device. | N/A |
|  | These procedures address the production and handling of media containing the MEDICAL DEVICE SOFTWARE including as appropriate: |  | N/A |
|  | - replication - media labelling - packaging - protection - storage - delivery | See above. | N/A |
| 6 | SOFTWARE MAINTENANCE PROCESS |  | — |
| 6.1 | [A, B, C] The MANUFACTURER establishes a software maintenance plan (or plans) for conducting the ACTIVITIES and TASKS of the maintenance PROCESS | Refer to Cl 5.1.15 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The plan addresses the following: |  | P |
|  | - procedures for: - receiving - documenting - evaluating - resolving - tracking feedback arising after release of the MEDICAL DEVICE SOFTWARE | Refer to Software Maintenance Guidelines (CUI-731.004) | P |
|  | - criteria for determining whether feedback is considered to be a problem | Refer to Cl 5.1.15 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - use of the software RISK MANAGEMENT PROCESS | Refer to Cl 5.1.15 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - use of the software problem resolution PROCESS for analysing and resolving problems arising after release of the MEDICAL DEVICE SOFTWARE | Refer to Cl 5.1.15 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 11 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - use of the software configuration management PROCESS for managing modifications to the existing SOFTWARE SYSTEM | Refer to Cl 5.1.15 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 10 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - procedures to EVALUATE and implement, for SOUP: |  |  |
|  | - upgrades - bug fixes - patches - obsolescence | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 6.2 | Problem and modification analysis |  | — |
| 6.2.1 | Document and EVALUATE feedback |  | P |
| 6.2.1.1 | [A, B, C] The MANUFACTURER monitors feedback on MEDICAL DEVICE SOFTWARE released for intended use | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.2.1.2 | [A, B, C] Feedback is documented and EVALUATED to determine whether a problem exists in a released MEDICAL DEVICE SOFTWARE | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
|  | Any such problem is recorded as a PROBLEM REPORT | See above. | P |
|  | PROBLEM REPORTS includes actual or potential adverse events, and deviations from specifications | See above. | P |
| 6.2.1.3 | [A, B, C] Each PROBLEM REPORT is EVALUATED to determine how it affects the SAFETY of MEDICAL DEVICE SOFTWARE released for intended use and whether a change to that SOFTWARE is needed to address the problem | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.2.2 | [A, B, C] The MANUFACTURER uses the software problem resolution PROCESS to address PROBLEM REPORTS | Refer to Cl 5.1.13 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.2.3 | [A, B, C] In addition to the analysis required by Clause 9, the MANUFACTURER analyses each CHANGE REQUEST for its effect on the organization, MEDICAL DEVICE SOFTWARE released for intended use, and SYSTEMS with which it interfaces | Refer to Cl 5.1.9.3 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.2.4 | [A, B, C] The MANUFACTURER EVALUATES and approves CHANGE REQUESTS which modify released MEDICAL DEVICE SOFTWARE | Refer to Cl 5.1.9.3 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.2.5 | [A, B, C] The MANUFACTURER identifies the approved CHANGE REQUESTS that affect released MEDICAL DEVICE SOFTWARE | Refer to Cl 5.1.9.3 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
|  | As required by local regulation, the MANUFACTURER informs users and regulators about: |  | P |
|  | - any problem in released MEDICAL DEVICE SOFTWARE and the consequences of continued unchanged use | See above. | P |
|  | - the nature of any available changes to released MEDICAL DEVICE SOFTWARE and how to obtain and install the changes | See above. | P |
| 6.3 | Modification implementation |  | — |
| 6.3.1 | [A, B, C] The MANUFACTURER identifies and perform any Clause 5 ACTIVITIES that need to be repeated as a result of the modification | Refer to Cl 5.1.9.3 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 6.3.2 | [A, B, C] The MANUFACTURER releases modifications according to 5.8 | See above. | P |
| 7 | SOFTWARE RISK MANAGEMENT PROCESS |  | — |
| 7.1 | Analysis of software contributing to hazardous situations |  | — |
| 7.1.1 | [B, C] The MANUFACTURER identifies SOFTWARE ITEMS that could contribute to a hazardous situation identified in the MEDICAL DEVICE RISK ANALYSIS ACTIVITY of ISO 14971 | Refer to Cl 5.1.7 of Software Development Plan [CRD-0601-PP-SDP-01] and Cl 5.3 of Software Design Specification (CRD-0601-DI-SDS-01) | P |
| 7.1.2 | [B, C] The MANUFACTURER identifies potential causes of the SOFTWARE ITEM identified above contributing to a hazardous situation | Refer to Cl 5.1.7 of Software Development Plan [CRD-0601-PP-SDP-01] and Cl 5.3 of Software Design Specification (CRD-0601-DI-SDS-01) | P |
|  | The MANUFACTURER considers potential causes including, as appropriate: |  | — |
|  | - incorrect or incomplete specification of functionality | See above. | P |
|  | - software defects in the identified SOFTWARE ITEM functionality | See above. | P |
|  | - failure or unexpected results from SOUP | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Guidelines (CUI-731.001) | P |
|  | - hardware failures or other software defects that could result in unpredictable software operation | See above. | P |
|  | - reasonably foreseeable misuse | See above. | P |
| 7.1.3 | [B, C] If failure or unexpected results from SOUP is a potential cause of the SOFTWARE ITEM contributing to a hazardous situation, the MANUFACTURER EVALUATES as a minimum any ANOMALY list published by the supplier of the SOUP item relevant to the VERSION of the SOUP item used in the MEDICAL DEVICE to determine if any of the known ANOMALIES result in a sequence of events that could result in a hazardous situation | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Guidelines (CUI-731.001) | P |
| 7.1.4 | [B, C] The MANUFACTURER documents in the RISK MANAGEMENT FILE potential causes of the SOFTWARE ITEM contributing to a hazardous situation | Refer to Cl 5.1.7 of Software Development Plan [CRD-0601-PP-SDP-01] and Cl 5.3 of Software Design Specification (CRD-0601-DI-SDS-01) | P |
| 7.2 | RISK CONTROL measures |  | — |
| 7.2.1 | [B, C] For each case documented in the RISK MANAGEMENT FILE where a SOFTWARE ITEM could contribute to a HAZARDOUS SITUATION, the MANUFACTURER defines and documents RISK CONTROL measures in accordance with ISO 14971 | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 7.2.2 | [B, C] If a RISK CONTROL measure is implemented as part of the functions of a SOFTWARE ITEM, the MANUFACTURER: |  | P |
|  | - includes the RISK CONTROL measure in the software requirements | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - assigns to each SOFTWARE ITEM that contributes to the implementation of a RISK CONTROL measure a software safety class based on the RISK that the RISK CONTROL measure is controlling | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - develops the SOFTWARE ITEM in accordance with Clause 5 | Refer to Cl 5.3 of Software Design Specification (CRD-0601-DI-SDS-01) | P |
| 7.3 | VERIFICATION of RISK CONTROL measures |  | — |
| 7.3.1 | [B, C] The implementation of each RISK CONTROL measure documented in 7.2 is VERIFIED, and this VERIFICATION is documented | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | The MANUFACTURER reviewers the RISK CONTROL measure and determines if it could result in a new HAZARDOUS SITUATION | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
| 7.3.3 | [B, C] The MANUFACTURER documents TRACEABILITY of software HAZARDS as appropriate: |  | P |
|  | - from the hazardous situation to the SOFTWARE ITEM | Refer to Cl 7 of Software Development File (CRD-0601-VV-SDF-01) | P |
|  | - from the SOFTWARE ITEM to the specific software cause | See above. | P |
|  | - from the software cause to the RISK CONTROL measure | See above. | P |
|  | - from the RISK CONTROL measure to the VERIFICATION of the RISK CONTROL measure | See above. | P |
| 7.4 | RISK MANAGEMENT of software changes |  | — |
| 7.4.1 | [A, B, C] The MANUFACTURER analyses changes to the MEDICAL DEVICE SOFTWARE (including SOUP) to determine whether: |  | P |
|  | - additional potential causes are introduced contributing to a hazardous situation | Refer to Risk Management Report (CRD-0601-VV-RMR-01) | P |
|  | - additional software RISK CONTROL measures are required | See above. | P |
| 7.4.2 | [B, C] The MANUFACTURER analyses changes to the software, including changes to SOUP, to determine whether the software modification could interfere with existing RISK CONTROL measures | Refer to Risk Management Report (CRD-0601-VV-RMR-01) | P |
| 7.4.3 | [B, C] The MANUFACTURER performs relevant RISK MANAGEMENT ACTIVITIES defined in 7.1, 7.2 and 7.3 based on these analyses | Refer to Risk Management Report (CRD-0601-VV-RMR-01) | P |
| 8 | SOFTWARE CONFIGURATION MANAGEMENT PROCESS |  | — |
| 8.1 | Configuration identification |  | — |
| 8.1.1 | [A, B, C] The MANUFACTURER establishes a scheme for the unique identification of CONFIGURATION ITEMS and their VERSIONS to be controlled according to the development and configuration planning specified in 5.1 | Refer to Cl 5.1.9 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 8.1.2 | [A, B, C] For each SOUP CONFIGURATION ITEM being used, including standard libraries, the MANUFACTURER documents: |  |  |
|  | - the title | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Guidelines (CUI-731.001) | P |
|  | - the MANUFACTURER | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Guidelines (CUI-731.001) | P |
|  | - the unique SOUP designator | Refer to Cl 5.1.12 of Software Development Plan (CRD-0601-PP-SDP-01) and Cl 9 of Software Development Guidelines (CUI-731.001) | P |
| 8.1.3 | [A, B, C] The MANUFACTURER documents the set of CONFIGURATION ITEMS and their VERSIONS that comprise the SOFTWARE SYSTEM configuration | Refer to Cl 5.1.9 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 8.2 | Change control |  | — |
| 8.2.1 | [A, B, C] The MANUFACTURER changes CONFIGURATION ITEMS identified to be controlled according to 8.1 only in response to an approved CHANGE REQUEST | Refer to Cl 5.1.9.3 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 8.2.2 | [A, B, C] The MANUFACTURER implements the change as specified in the CHANGE REQUEST | Refer to Cl 5.1.9.3 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The MANUFACTURER identifies and performs any ACTIVITY that needs to be repeated as a result of the change, including changes to the software safety classification of SOFTWARE SYSTEMS and SOFTWARE ITEMS | See above. | P |
| 8.2.3 | [A, B, C] The MANUFACTURER verifies the change, including repeating any VERIFICATION that has been invalidated by the change and taking into account 5.7.3 and 9.7 | Refer to Cl 5.1.9.3 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 8.2.4 | [A, B, C] The MANUFACTURER maintains records of the relationships and dependencies between: |  | P |
|  | - CHANGE REQUEST | Refer to Design Change Procedure (CUP-737) | P |
|  | - relevant PROBLEM REPORT | See above. | P |
|  | - approval of the CHANGE REQUEST | See above. | P |
| 8.3 | [A, B, C] The MANUFACTURER retains retrievable records of the history of controlled CONFIGURATION ITEMS including SYSTEM configuration | Refer to Cl 5.1.9.3 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 9 | SOFTWARE PROBLEM RESOLUTION PROCESS |  | — |
| 9.1 | [A, B, C] The MANUFACTURER prepares a PROBLEM REPORT for each problem detected in the MEDICAL DEVICE SOFTWARE | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | PROBLEM REPORTS include a statement of criticality (for example, effect on performance, SAFETY, or SECURITY) as well as other information that may aid in the resolution of the problem (for example, devices affected, supported accessories affected) | See above. | P |
| 9.2 | [A, B, C] The MANUFACTURER: |  | P |
|  | - investigates the problem and if possible identify the causes | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - EVALUATES the problem’s relevance to SAFETY using the software RISK MANAGEMENT PROCESS | See above. | P |
|  | - documents the outcome of the investigation and evaluation | See above. | P |
|  | - creates a CHANGE REQUEST(S) for actions needed to correct the problem, or document the rationale for taking no action | See above. | P |
| 9.3 | [A, B, C] The MANUFACTURER advises relevant parties of the existence of the problem, as appropriate | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 9.4 | [A, B, C] The MANUFACTURER approves and implements all CHANGE REQUESTS, observing the requirements of the change control PROCESS | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 9.5 | [A, B, C] The MANUFACTURER maintains records of PROBLEM REPORTS and their resolution including their VERIFICATION | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | The MANUFACTURER updates the RISK MANAGEMENT FILE as appropriate | Refer to Risk Management Plan [CRD-0601-PP-RMP-01] and Cl 5.1.7 of Software Development Plan [CRD-0601-PP-SDP-01]. | P |
| 9.6 | [A, B, C] The MANUFACTURER performs analysis to detect trends in PROBLEM REPORTS | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
| 9.7 | [A, B, C] The MANUFACTURER verifies resolutions to determine whether: |  | P |
|  | - problem has been resolved and the PROBLEM REPORT has been closed | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - adverse trends have been reversed | See above. | P |
|  | - CHANGE REQUESTS have been implemented in the appropriate MEDICAL DEVICE SOFTWARE and ACTIVITIES | See above. | P |
|  | - additional problems have been introduced | See above. | P |
| 9.8 | [A, B, C] When testing, retesting or REGRESSION TESTING SOFTWARE ITEMS and SYSTEMS following a change, the MANUFACTURER includes in the test documentation: |  | P |
|  | - test results | Refer to Cl 5.1.13 of Software Development Plan (CRD-0601-PP-SDP-01) | P |
|  | - ANOMALIES found | See above. | P |
|  | - the VERSION of software tested | See above. | P |
|  | - relevant hardware and software test configurations | See above. | P |
|  | - relevant test tools | See above. | P |
|  | - date tested | See above. | P |
|  | - identification of the tester | See above. | P |

- Appendix 2. Traceability Matrix

- SRS – SAD – SDS

| SRS | SAD | SDS |
| --- | --- | --- |
| FUN-REQ-01 | SAD-EA-01, SAD-EA-02, SAD-PD-04, SAD-EA-03, SAD-EA-04, SAD-EA-05 | SDS-MA-021,SDS-MA-022,SDS-MA-084,SDS-MA-024,SDS-MA-025,SDS-MA-015,SDS-MA-029,SDS-MA-083,SDS-MA-042,SDS-MA-068,SDS-MA-069 |
| FUN-REQ-02 | SAD-SE-01, SAD-SE-02, SAD-SE-03, SAD-CT-01, SAD-CA-02, SAD-HV-01, SAD-HV-02, SAD-HV-04, SAD-EA-02, SAD-UI-01 | SDS-MA-041,SDS-MA-043,SDS-MA-066,SDS-MA-067,SDS-MA-094,SDS-TP-008,SDS-TP-009,SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-TP-001,SDS-TP-002,SDS-TP-003,SDS-TP-007,SDS-TP-004,SDS-TP-006,SDS-MA-024,SDS-MA-025,SDS-MA-086,SDS-MA-087,SDS-MA-088 |
| FUN-REQ-03 | SAD-UI-01, SAD-UI-02, SAD-UI-06, SAD-UI-07, SAD-UI-09 | SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-044,SDS-MA-031,SDS-MA-027,SDS-MA-028,SDS-MA-089 |
| FUN-REQ-04 | SAD-BM-01, SAD-BM-03, SAD-BM-04, SAD-DV-02, SAD-DV-04, SAD-PD-01, SAD-PD-04, SAD-UI-09 | SDS-MA-020,SDS-MA-018,SDS-MA-01,SDS-MA-072,SDS-MA-004,SDS-MA-026,SDS-MA-073,SDS-MA-011,SDS-MA-071,SDS-MA-015,SDS-MA-029,SDS-MA-083, SDS-MA-089 |
| MNT-REQ-01 | SAD-BM-01, SAD-BM-02, SAD-BM-03, SAD-BM-04, SAD-CA-02, SAD-CA-03, SAD-CM-01, SAD-CT-01, SAD-HV-01, SAD-HV-03, SAD-EE-01, SAD-ST-01, SAD-UI-09, SAD-UM-02 | SDS-MA-020,SDS-MA-018,SDS-MA-072,SDS-MA-019,SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-MA-079,SDS-MA-080,SDS-MP-013,SDS-MP-014,SDS-MP-015,SDS-TP-008,SDS-TP-009,SDS-TP-001,SDS-TP-002,SDS-TP-003,SDS-TP-007,SDS-TP-005,SDS-MP-016,SDS-MP-017,SDS-MP-018,SDS-MP-019,SDS-MP-020,SDS-MP-021,SDS-MP-022,SDS-MP-023,SDS-MA-058,SDS-MA-059,SDS-MA-089,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SAD-UM-02 |
| MNT-REQ-02 | SAD-BM-01, SAD-BM-02, SAD-BM-03, SAD-BM-04, SAD-CA-02, SAD-CA-03, SAD-CM-01, SAD-CT-01, SAD-HV-01, SAD-HV-03, SAD-EE-01, SAD-PD-01, SAD-PD-02, SAD-PD-03, SAD-ST-02, SAD-UI-01, SAD-UM-02, SAD-ST-02 | SDS-MA-020,SDS-MA-017,SDS-MA-018,SDS-MA-072,SDS-MA-019,SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-MA-079,SDS-MA-080,SDS-MP-013,SDS-MP-014,SDS-MP-015,SDS-TP-008,SDS-TP-009,SDS-TP-001,SDS-TP-002,SDS-TP-003,SDS-TP-007,SDS-TP-005,SDS-MP-016,SDS-MP-017,SDS-MP-018,SDS-MP-019,SDS-MP-020,SDS-MP-021,SDS-MP-022,SDS-MP-023,SDS-MA-011,SDS-MA-029,SDS-MA-071,SDS-MA-012,SDS-MA-013,SDS-MA-014,SDS-MA-083,SDS-MA-060,SDS-MA-061,SDS-MA-062,SDS-MA-063,SDS-MA-074,SDS-MA-090,SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SDS-MP-030 |
| MNT-REQ-03 | SAD-ST-03, SAD-BM-01, SAD-BM-02, SAD-BM-03, SAD-BM-04, SAD-CA-02, SAD-CA-03, SAD-PD-01, SAD-PD-02, SAD-PD-03, SAD-CM-01, SAD-EE-01, SAD-HV-01, SAD-HV-03, SAD-PD-01, SAD-PD-03, SAD-CT-01, SAD-UI-01, SAD-UI-09, SAD-UM-02 | SDS-MA-064,SDS-MA-065, SDS-MA-020,SDS-MA-017,SDS-MA-018,SDS-MA-019,SDS-MA-072,SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-MA-079,SDS-MA-080,SDS-MA-011,SDS-MA-029,SDS-MA-071,SDS-MA-014,SDS-MA-083,SDS-MP-013,SDS-MP-014,SDS-MP-015,SDS-MP-016,SDS-MP-017,SDS-MP-018,SDS-MP-019,SDS-MP-020,SDS-MP-021,SDS-MP-022,SDS-MP-023,SDS-TP-001,SDS-TP-002,SDS-TP-003,SDS-TP-007,SDS-TP-005,SDS-TP-008,SDS-TP-009,SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-089,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SDS-MP-030 |
| MNT-REQ-04 | SAD-DM-01, SAD-DM-02, SAD-DM-03, SAD-DV-07, SAD-CA-01, SAD-CA-04 | SDS-MA-005,SDS-MA-006,SDS-MA-007,SDS-MA-023,SDS-MA-048,SDS-MA-045,SDS-MA-075,SDS-MA-076,SDS-MA-070,SDS-MA-081,SDS-MA-082 |
| NFN-REQ-01 | SAD-BM-03, SAD-BM-04, SAD-DV-03, SAD-UI-01, SAD-UM-02 | SDS-MA-018,SDS-MA-072,SDS-MA-01,SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SAD-UM-02 |
| NFN-REQ-02 | SAD-UI-01, SAD-DV-05, SAD-ST-01 | SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-030,SDS-MA-058,SDS-MA-059 |
| NFN-REQ-03 | SAD-DM-01, SAD-DV-07, SAD-DV-08 | SDS-MA-005,SDS-MA-006,SDS-MA-007,SDS-MA-045,SDS-MA-047 |
| NFN-REQ-04 | SAD-CA-01, SAD-CA-05, SAD-DV-09, SAD-DV-10, SAD-DV-12 | SDS-MA-075,SDS-MA-076,SDS-MA-091,SDS-MA-092,SDS-MA-093,SDS-MA-049,SDS-MA-050,SDS-MA-051,SDS-MA-052,SDS-MA-053,SDS-MA-054,SDS-MA-057,SDS-MA-055,SDS-MA-056,SDS-MA-001 |
| NFN-REQ-05 | SAD-DV-01, SAD-DV-06, SAD-DV-12, SAD-PM-01, SAD-PM-02, SAD-PM-03, SAD-PM-04, SAD-PM-05, SAD-CA-03, SAD-CM-01, SAD-UM-01, SAD-UM-02, SAD-EE-01 | SDS-MA-002,SDS-MA-003,SDS-MA-008,SDS-MA-035,SDS-MA-036,SDS-MA-037,SDS-MA-038,SDS-MP-002,SDS-MP-003,SDS-MP-004,SDS-MP-005,SDS-MP-006,SDS-MP-007,SDS-MP-008,SDS-MP-001,SDS-MP-009,SDS-MP-010,SDS-MP-011,SDS-MP-012,SDS-MP-013,SDS-MP-014,SDS-MP-015,SDS-MP-026,SDS-MP-027,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SDS-MP-030,SDS-MP-016,SDS-MP-017,SDS-MP-018,SDS-MP-019,SDS-MP-020,SDS-MP-021,SDS-MP-022,SDS-MP-023 |
| SFT-REQ-01 | SAD-CA-02, SAD-SE-02, SAD-HV-03, SAD-DV-02, SAD-PD-04 | SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-MA-043,SDS-TP-005,SDS-MA-004,SDS-MA-015,SDS-MA-029,SDS-MA-083 |
| SFT-REQ-02 | SAD-CA-02, SAD-HV-01, SAD-HV-03, SAD-UI-01, SAD-SE-03 | SDS-MA-070,SDS-MA-077,SDS-MA-078,SDS-TP-001,SDS-TP-002,SDS-TP-003,SDS-TP-007,SDS-TP-005,SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-066,SDS-MA-067,SDS-MA-094 |
| UI-REQ-01 | SAD-UI-01, SAD-UI-03, SAD-UI-05, SAD-UI-08, SAD-UM-02, SAD-PD-01 | SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-040,SDS-MA-032,SDS-MA-033,SDS-MA-034,SDS-MA-016,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SDS-MP-030,SDS-MA-011,SDS-MA-029,SDS-MA-071 |
| UI-REQ-02 | SAD-UI-01, SAD-UI-04, SAD-UM-02 | SDS-MA-086,SDS-MA-087,SDS-MA-088,SDS-MA-039,SDS-MP-024,SDS-MP-025,SDS-MP-028,SDS-MP-029,SDS-MP-030 |

- SAD – SDS

| SAD | SDS |
| --- | --- |
| SAD-EA-01 | SDS-MA-021, SDS-MA-022, SDS-MA-084 |
| SAD-EA-02 | SDS-MA-024, SDS-MA-025 |
| SAD-EA-03 | SDS-MA-042 |
| SAD-EA-04 | SDS-MA-068 |
| SAD-EA-05 | SDS-MA-069 |
| SAD-SE-01 | SDS-MA-041 |
| SAD-SE-02 | SDS-MA-043 |
| SAD-SE-03 | SDS-MA-066, SDS-MA-067, SDS-MA-094 |
| SAD-PD-01 | SDS-MA-011, SDS-MA-029, SDS-MA-071 |
| SAD-PD-02 | SDS-MA-012, SDS-MA-013 |
| SAD-PD-03 | SDS-MA-014, SDS-MA-083 |
| SAD-PD-04 | SDS-MA-015, SDS-MA-029, SDS-MA-083 |
| SAD-BM-01 | SDS-MA-020 |
| SAD-BM-02 | SDS-MA-017 |
| SAD-BM-03 | SDS-MA-018, SDS-MA-072 |
| SAD-BM-04 | SDS-MA-019, SDS-MA-072 |
| SAD-ST-01 | SDS-MA-058, SDS-MA-059 |
| SAD-ST-02 | SDS-MA-060, SDS-MA-061, SDS-MA-062, SDS-MA-063, SDS-MA-074, SDS-MA-090 |
| SAD-ST-03 | SDS-MA-064, SDS-MA-065 |
| SAD-UI-01 | SDS-MA-086, SDS-MA-087, SDS-MA-088 |
| SAD-UI-02 | SDS-MA-044 |
| SAD-UI-03 | SDS-MA-040 |
| SAD-UI-04 | SDS-MA-039 |
| SAD-UI-05 | SDS-MA-032, SDS-MA-033, SDS-MA-034 |
| SAD-UI-06 | SDS-MA-031 |
| SAD-UI-07 | SDS-MA-027, SDS-MA-028 |
| SAD-UI-08 | SDS-MA-016 |
| SAD-UI-09 | SDS-MA-089 |
| SAD-DM-01 | SDS-MA-005, SDS-MA-006, SDS-MA-007 |
| SAD-DM-02 | SDS-MA-023 |
| SAD-DM-03 | SDS-MA-048 |
| SAD-DV-01 | SDS-MA-002, SDS-MA-003, SDS-MA-008 |
| SAD-DV-02 | SDS-MA-004 |
| SAD-DV-03 | SDS-MA-009, SDS-MA-010 |
| SAD-DV-04 | SDS-MA-026, SDS-MA-073 |
| SAD-DV-05 | SDS-MA-030 |
| SAD-DV-06 | SDS-MA-035, SDS-MA-036, SDS-MA-037, SDS-MA-038 |
| SAD-DV-07 | SDS-MA-045 |
| SAD-DV-08 | SDS-MA-047 |
| SAD-DV-09 | SDS-MA-049, SDS-MA-050, SDS-MA-051, SDS-MA-052, SDS-MA-053, SDS-MA-054, SDS-MA-057 |
| SAD-DV-10 | SDS-MA-055, SDS-MA-056 |
| SAD-DV-11 | SDS-MA-046, SDS-MA-085 |
| SAD-DV-12 | SDS-MA-001 |
| SAD-CA-01 | SDS-MA-075, SDS-MA-076 |
| SAD-CA-02 | SDS-MA-070, SDS-MA-077, SDS-MA-078 |
| SAD-CA-03 | SDS-MA-070, SDS-MA-079, SDS-MA-080 |
| SAD-CA-04 | SDS-MA-070, SDS-MA-081, SDS-MA-082 |
| SAD-CA-05 | SDS-MA-091, SDS-MA-092, SDS-MA-093 |
| SAD-CM-01 | SDS-MP-013, SDS-MP-014, SDS-MP-015 |
| SAD-UM-01 | SDS-MP-026, SDS-MP-027 |
| SAD-UM-02 | SDS-MP-024, SDS-MP-025, SDS-MP-028, SDS-MP-029 |
| SAD-EE-01 | SDS-MP-016, SDS-MP-017, SDS-MP-018, SDS-MP-019, SDS-MP-020, SDS-MP-021, SDS-MP-022, SDS-MP-023 |
| SAD-PM-01 | SDS-MP-002, SDS-MP-003, SDS-MP-004, SDS-MP-005, SDS-MP-006, SDS-MP-007, SDS-MP-008 |
| SAD-PM-02 | SDS-MP-001 |
| SAD-PM-03 | SDS-MP-009 |
| SAD-PM-04 | SDS-MP-010 |
| SAD-PM-05 | SDS-MP-011, SDS-MP-012 |
| SAD-CT-01 | SDS-TP-008, SDS-TP-009 |
| SAD-HV-01 | SDS-TP-001, SDS-TP-002, SDS-TP-003, SDS-TP-007 |
| SAD-HV-02 | SDS-TP-004 |
| SAD-HV-03 | SDS-TP-005 |
| SAD-HV-04 | SDS-TP-006 |

- SRS – SST

| SRS ID | TEST ID | Test Report Document Name |
| --- | --- | --- |
| FUN-REQ-01 | SST-01 | [CRD-0601-0434-260105-001] ECG Measurement and Analysis System Test Report |
| FUN-REQ-02 | SST-02 | [CRD-0601-0434-260108-001] Shock Energy Delivery (External Discharge) System Test Report |
| FUN-REQ-03 | SST-03-01 | [CRD-0601-0569-260112-001] CPR Guidance-1 System Test Report |
|  | SST-03-02 | [CRD-0601-0569-260112-002] CPR Guidance-2 System Test Report |
|  | SST-03-03 | [CRD-0601-0569-260112-003] CPR Guidance-3 (Pads Attachment Guide - detailed mode) System Test Report |
|  | SST-03-04 | [CRD-0601-0569-260112-004] CPR Guidance-3 (Pads Attachment Guide - shortened mode) System Test Report |
|  | SST-03-05 | [CRD-0601-0569-260112-005] AHA Guidelines Compliance System Test Report |
| FUN-REQ-04 | SST-04 | [CRD-0601-0434-260115-001] Device Status Monitoring System Test Report |
| UI-REQ-01 | SST-05-01 | [CRD-0601-0569-260116-001] Pads Management (Defibrillator Pads Connector) System Test Report |
|  | SST-05-02 | [CRD-0601-0569-260116-002] Pads Management (Status and Attachment) System Test Report |
| UI-REQ-02 | SST-06 | [CRD-0601-0434-260120-001] Device Operation Status Guidance System Test Report |
| SFT-REQ-01 | SST-07 | [CRD-0601-0569-260122-001] Shock Energy Delivery (Internal Discharge) System Test Report |
| SFT-REQ-02 | SST-08-01 | [CRD-0601-0569-260123-001] Real-time Monitoring System for High Voltage Capacitor Charging Voltage Test Report |
|  | SST-08-02 | [CRD-0601-0569-260123-002] High Voltage Capacitor Overcharge Monitoring System Test Report |
| MNT-REQ-01 | SST-09 | [CRD-0601-0434-260127-001] POST(Power-On Self-Test) System Test |
| MNT-REQ-02 | SST-10-01 | [CRD-0601-0569-260129-001] PST(Periodic Self-Test) System Test |
|  | SST-10-02 | [CRD-0601-0569-260129-002] PST(Periodic Self-Test) Error Detection System Test Report |
| MNT-REQ-03 | SST-11 | [CRD-0601-0434-260202-001] BIT(Battert Insertion Test) System Test |
| MNT-REQ-04 | SST-12 | [CRD-0601-0434-260202-002] Patient Treatment Data Storage and Export System Test Report |
| NFN-REQ-01 | SST-13 | [CRD-0601-0434-260202-003] Device Unavailability Guidance System Test Report |
| NFN-REQ-02 | SST-14 | [CRD-0601-0569-260209-001] Error CPR Guidance System Test |
| NFN-REQ-03 | SST-15 | [CRD-0601-0569-260210-001] CPR Guidance Settings System Test |
| NFN-REQ-04 | SST-16-01 | [CRD-0601-0569-260211-001] Factory Mode Settings System Test |
|  | SST-16-02 | [CRD-0601-0569-260211-002] Factory Mode Settings Change System Test |
|  | SST-16-03 | [CRD-0601-0569-260211-003] Factory Mode Settings Change via IrDA and SD Card System Test Report |
| NFN-REQ-05 | SST-17 | [CRD-0601-0434-260213-001] AED Standby Mode Maintenance System Test Report |

- SAD –SIT

| SAD | SIT(AD) | Test Report Document Name |
| --- | --- | --- |
| SAD-EA-01 | SIT-EA-01 | [CRD-0601-0434-250804-001] ECG Data Collection Test Report |
|  | SIT-EA-02 | [CRD-0601-0434-250804-002] ECG Data Filtering Test Report |
| SAD-EA-02 | SIT-EA-03 | [CRD-0601-0434-250804-003] ECG Data Analyzing Test Report |
| SAD-EA-03 | SIT-EA-03 | [CRD-0601-0434-250804-003] ECG Data Analyzing Test Report |
| SAD-EA-04 | SIT-EA-03 | [CRD-0601-0434-250804-003] ECG Data Analyzing Test Report |
| SAD-EA-05 | SIT-EA-03 | [CRD-0601-0434-250804-003] ECG Data Analyzing Test Report |
| SAD-SE-01 | SIT-SE-01 | [CRD-0601-0434-250807-001] Patient Mode–Based Energy Charging Start Test Report |
| SAD-SE-02 | SIT-SE-02 | [CRD-0601-0434-250807-002] ECG Analysis–Based Energy Delivery Decision and User Guidance Test Report |
| SAD-SE-03 | SIT-SE-03-01 | [CRD-0601-0434-250807-003] Charge/Discharge Step-wise Command Transmitting Test Report |
|  | SIT-SE-03-02 | [CRD-0601-0434-250807-004] Therapy B/D Charging Status Monitoring Test Report |
|  | SIT-SE-03-03 | [CRD-0601-0434-250807-005] Therapy B/D Communication Timeout Test Report |
| SAD-PD-01 | SIT-PD-01 | [CRD-0601-0434-250811-001] Pads Connector Connection Status Test Report |
| SAD-PD-02 | SIT-PD-02 | [CRD-0601-0434-250811-002] Pads Internal Memory Verification Test Report |
| SAD-PD-03 | SIT-PD-03 | [CRD-0601-0434-250811-003] Pad Storage Condition Assessment Verification Test Report |
| SAD-PD-04 | SIT-PD-04 | [CRD-0601-0434-250811-004] Pads Attachment Status Realtime Monitoring Test Report |
| SAD-BM-01 | SIT-BM-01 | [CRD-0601-0569-250814-001] Battery Level Classification and State Detection Integration Test Report |
| SAD-BM-02 | SIT-BM-02 | [CRD-0601-0569-250814-002] Residual Battery Capacity Calculation Test Report |
| SAD-BM-03 | SIT-BM-03 | [CRD-0601-0569-250814-003] Low Battery Status Detection Test Report |
| SAD-BM-04 | SIT-BM-04 | [CRD-0601-0569-250814-004] Battery Abnormality Diagnosis Verification Test Report |
| SAD-ST-01 | SIT-ST-01 | [CRD-0601-0569-250818-001] POST Test Report |
| SAD-ST-02 | SIT-ST-02 | [CRD-0601-0569-250818-002] POST Cycle Execution Test Report |
| SAD-ST-03 | SIT-ST-03 | [CRD-0601-0569-250818-003] BIT Test Report |
| SAD-UI-01 | SIT-UI-01-01 | [CRD-0601-0569-250821-001] Main ARM UI Output Integration Test Report |
|  | SIT-UI-01-02 | [CRD-0601-0569-250821-002] Device Status Notification Test (Low Battery Status) |
|  | SIT-UI-01-03 | [CRD-0601-0569-250821-003] Device Status Notification Test (System Error Status) |
| SAD-UI-02 | SIT-UI-02-01 | [CRD-0601-0569-250822-001] CPR Guide Mode Entry Integration Test Report |
|  | SIT-UI-02-02 | [CRD-0601-0569-250822-002] Post-Defibrillation CPR Guidance in CPR Guide Mode Integration Test Report |
|  | SIT-UI-02-03 | [CRD-0601-0569-250822-003] CPR Guidance Transition After Internal Discharge Integration Test Report |
| SAD-UI-03 | SIT-UI-03 | [CRD-0601-0569-250825-001] Pad Attachment Instructions Test Report |
| SAD-UI-04 | SIT-UI-04 | [CRD-0601-0569-250825-002] Device Startup Guidance Test Report |
| SAD-UI-05 | SIT-UI-05 | [CRD-0601-0569-250825-003] Patient Mode Determination Test Report |
| SAD-UI-06 | SIT-UI-06 | [CRD-0601-0569-250825-004] CPR Activity Detection Test Report |
| SAD-UI-07 | SIT-UI-07-01 | [CRD-0601-0569-250828-001] Voice Output-based Integrated CPR Guidance Test Report |
|  | SIT-UI-07-02 | [CRD-0601-0569-250828-002] Operation Mode-based Integrated CPR Guidance Test Report |
| SAD-UI-08 | SIT-UI-08 | [CRD-0601-0569-250828-003] Pads Attachment Voice Prompt Setup Test Report |
| SAD-UI-09 | SIT-UI-09 | [CRD-0601-0569-250828-004] Main ARM UI Input Integration Test Report |
| SAD-DM-01 | SIT-DM-01 | [CRD-0601-0434-250904-001] Norflash Memory Data Storage Test Report |
| SAD-DM-02 | SIT-DM-02 | [CRD-0601-0434-250904-002] ECG Data Storage Test Report |
| SAD-DM-03 | SIT-DM-03 | [CRD-0601-0434-250904-003] Treatment Record Extraction Test Report |
| SAD-DV-01 | SIT-DV-01 | [CRD-0601-0569-250911-001] Wake-up Status Check and Initialization Test Report |
| SAD-DV-02 | SIT-DV-02 | [CRD-0601-0569-250911-002] Real-time Power-off Detection Monitoring Test Report |
| SAD-DV-03 | SIT-DV-03-01 | [CRD-0601-0569-250911-003] System Error Detection and Error Type Classification Test Report |
|  | SIT-DV-03-02 | [CRD-0601-0569-250911-004] Error Code Format and Range Validation with Retry Handling Test Report |
|  | SIT-DV-03-03 | [CRD-0601-0569-250911-005] User Notification and Power Shutdown Control for Persistent Error Conditions Test Report |
| SAD-DV-04 | SIT-DV-04 | [CRD-0601-0434-250918-001] Environment Noise Analysis and Voice Prompt Volume Control Test Report |
| SAD-DV-05 | SIT-DV-05 | [CRD-0601-0434-250918-002] CPR Guidance for System Error Conditions Test Report |
| SAD-DV-06 | SIT-DV-06 | [CRD-0601-0434-250918-003] System Mode Determination Test Report |
| SAD-DV-07 | SIT-DV-07 | [CRD-0601-0434-250925-001] Administrator Mode Entry Test Report |
| SAD-DV-08 | SIT-DV-08 | [CRD-0601-0434-250925-002] CPR Guidance Setup Update Test Report |
| SAD-DV-09 | SIT-DV-09-01 | [CRD-0601-0434-250925-003] Factory Mode Device Setup Test Report |
|  | SIT-DV-09-02 | [CRD-0601-0434-250925-004] Factory Mode Device Initialization Setup Test Report |
| SAD-DV-10 | SIT-DV-10 | [CRD-0601-0569-251002-001] Factory Mode SD Card-Based Device Setup Test Report |
| SAD-DV-11 | SIT-DV-11 | [CRD-0601-0569-251002-002] IrDA-Based Device Setup Change Test Report |
| SAD-DV-12 | SIT-DV-12 | [CRD-0601-0569-251002-003] Post-Power-on Idle State Monitoring Test Report |
| SAD-CA-01 | SIT-CA-01 | [CRD-0601-0434-251009-001] SD Card Read/Write Operations Test Report |
| SAD-CA-02 | SIT-CA-02 | [CRD-0601-0434-251009-002] Therapy PIC-High Voltage Module Communication Test Report |
| SAD-CA-03 | SIT-CA-03 | [CRD-0601-0434-251009-003] Main PIC–EEPROM Control Module Communication Test Report |
| SAD-CA-04 | SIT-CA-04 | [CRD-0601-0434-251009-004] IrDA Communication Test Report |
| SAD-CM-01 | SIT-CM-01 | [CRD-0601-0434-251010-001] Main PIC I2C Communication Test Report |
| SAD-UM-01 | SIT-UM-01 | [CRD-0601-0569-251016-001] Main PIC UI Input Integration Test Report |
| SAD-UM-02 | SIT-UM-02-01 | [CRD-0601-0569-251016-002] Main PIC UI Output Control (LED) Test Report |
|  | SIT-UM-02-02 | [CRD-0601-0569-251016-003] Main PIC UI Output Control (LCD) Test Report |
| SAD-EE-01 | SIT-EE-01 | [CRD-0601-0569-251023-001] Pads and Battery Information Test Report |
| SAD-PM-01 | SIT-PM-01 | [CRD-0601-0569-251023-002] Main ARM Power Activation Test Report |
| SAD-PM-02 | SIT-PM-02 | [CRD-0601-0569-251023-003] Main ARM Status Monitoring Test Report |
| SAD-PM-03 | SIT-PM-03 | [CRD-0601-0569-251023-004] Main ARM Power Deactivation Test Report |
| SAD-PM-04 | SIT-PM-04 | [CRD-0601-0569-251030-001] Low Power Standby Mode Transition Test Report |
| SAD-PM-05 | SIT-PM-05 | [CRD-0601-0569-251030-002] Main PIC RTC Control Test Report |
| SAD-CT-01 | SIT-CT-01 | [CRD-0601-0569-251030-003] Main ARM–Therapy PIC Integration Communication Test Report |
| SAD-HV-01 | SIT-HV-01 | [CRD-0601-0434-251105-001] High Voltage Module Energy Charging Control and Status Check Test Report |
| SAD-HV-02 | SIT-HV-02 | [CRD-0601-0434-251105-002] External Shock Energy Discharge Test Report |
| SAD-HV-03 | SIT-HV-03 | [CRD-0601-0434-251105-003] Internal Shock Energy Discharge Test Report |
| SAD-HV-04 | SIT-HV-04 | [CRD-0601-0434-251105-004] External Defibrillation Biphasic Waveform Integration Test Report |

- SDS-SUT

| SDS ID | TEST ID | Test Report Document Name |
| --- | --- | --- |
| SDS-MA-001 | SUT-MA-001 | CRD-0601-0434-241217-001_PwrSleepTask |
| SDS-MA-002 | SUT-MA-002 | CRD-0601-0434-241217-002_InitRecord |
|  | SUT-MA-003 | CRD-0601-0434-241217-003__System_Read |
|  | SUT-MA-004 | CRD-0601-0434-241217-004_Dev_Operation_Data_Read |
|  | SUT-MA-005 | CRD-0601-0434-241217-005_Dev_Config_Data_Read |
| SDS-MA-003 | SUT-MA-006 | CRD-0601-0434-241217-006_AED_Main |
| SDS-MA-004 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
|  | SUT-MA-008 | CRD-0601-0434-241217-008_Power_Off |
|  | SUT-MA-009 | CRD-0601-0434-241217-009_PwrKeyProcTask |
|  | SUT-MA-010 | CRD-0601-0434-241217-010_setPSTIntervalCounter |
| SDS-MA-005 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
|  | SUT-MA-011 | CRD-0601-0434-241217-011_EventRecord |
| SDS-MA-006 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
|  | SUT-MA-011 | CRD-0601-0434-241217-011_EventRecord |
|  | SUT-MA-012 | CRD-0601-0434-241217-012_ECGRecord |
| SDS-MA-007 | SUT-MA-013 | CRD-0601-0434-241217-013_Dev_Config_Data_UpDate |
| SDS-MA-008 | SUT-MA-014 | CRD-0601-0434-241217-014_ResetModeTask |
| SDS-MA-009 | SUT-MA-015 | CRD-0601-0434-241217-015_SystemMonitorTask |
| SDS-MA-010 | SUT-MA-016 | CRD-0601-0434-241217-016_SystemErrorMsg |
|  | SUT-MA-017 | CRD-0601-0434-241217-017_SystemErrorMsgduringCharging |
| SDS-MA-011 | SUT-MA-018 | CRD-0601-0434-241217-018_PadsConnectChkMsg |
|  | SUT-MA-019 | CRD-0601-0434-241217-019_PadsProcTask |
| SDS-MA-012 | SUT-MA-020 | CRD-0601-0434-241217-020_Pads_EEPROM_Use_Task |
| SDS-MA-013 | SUT-MA-021 | CRD-0601-0434-241217-021_Set_Pads_Exp_Flag |
| SDS-MA-014 | SUT-MA-022 | CRD-0601-0434-241217-022_PadsCheckTask |
| SDS-MA-015 | SUT-MA-038 | CRD-0601-0434-241218-006_AEDPadsOff |
|  | SUT-MA-042 | CRD-0601-0434-241218-010_CPRImpProcTask |
|  | SUT-MA-019 | CRD-0601-0434-241217-019_PadsProcTask |
| SDS-MA-016 | SUT-MA-023 | CRD-0601-0434-241217-022_PadsCheckTask |
| SDS-MA-017 | SUT-MA-025 | CRD-0601-0434-241217-025_BatteryLevelCalculation_withCounter |
|  | SUT-MA-024 | CRD-0601-0434-241217-024_RunningBatteryLevelCalculation |
| SDS-MA-018 | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
| SDS-MA-019 | SUT-MA-027 | CRD-0601-0434-241217-027_StopSystemWhenBatteryDepleted |
| SDS-MA-020 | SUT-MA-028 | CRD-0601-0434-241217-028_BatteryMonitorTask |
| SDS-MA-021 | SUT-MA-029 | CRD-0601-0434-241217-029_ECGAnalSegProc |
| SDS-MA-022 | SUT-MA-030 | CRD-0601-0434-241217-030_ECGFilterOut |
|  | SUT-MA-031 | CRD-0601-0434-241217-031_ECGPreProcTask |
| SDS-MA-023 | SUT-MA-012 | CRD-0601-0434-241217-012_ECGRecord |
| SDS-MA-024 | SUT-MA-032 | CRD-0601-0434-241217-032_ECGAnalyzeTask |
| SDS-MA-025 | SUT-MA-032 | CRD-0601-0434-241217-032_ECGAnalyzeTask |
| SDS-MA-026 | SUT-MA-033 | CRD-0601-0434-241218-001_AudioVolumeTask |
| SDS-MA-027 | SUT-MA-034 | CRD-0601-0434-241218-002_CPR_Selection_by_iBtn |
| SDS-MA-028 | SUT-MA-036 | CRD-0601-0434-241218-004_CPRChgTask |
|  | SUT-MA-035 | CRD-0601-0434-241218-003_CPRProcTask |
| SDS-MA-029 | SUT-MA-037 | CRD-0601-0434-241218-005_checkPadsStatusAfterCPR |
|  | SUT-MA-038 | CRD-0601-0434-241218-006_AEDPadsOff |
| SDS-MA-030 | SUT-MA-039 | CRD-0601-0434-241218-007_SysErrCPRChgTask |
|  | SUT-MA-041 | CRD-0601-0434-241218-009_SysErrCPRProcTask |
|  | SUT-MA-040 | CRD-0601-0434-241218-008_SysErrCPRProtocolTask |
| SDS-MA-031 | SUT-MA-042 | CRD-0601-0434-241218-010_CPRImpProcTask |
| SDS-MA-032 | SUT-MA-043 | CRD-0601-0434-241218-011_AdultPadsMsgTask |
| SDS-MA-033 | SUT-MA-044 | CRD-0601-0434-241218-012_InfantPadsMsgTask |
|  | SUT-MA-045 | CRD-0601-0434-241218-013_ModeChangeTask |
| SDS-MA-034 | SUT-MA-045 | CRD-0601-0434-241218-013_ModeChangeTask |
| SDS-MA-035 | SUT-MA-046 | CRD-0601-0434-241218-014_AED_mainStart |
|  | SUT-MA-047 | CRD-0601-0434-241218-015_FactoryModeChk |
|  | SUT-MA-048 | CRD-0601-0434-241218-016_FM_System_Field_Mode |
|  | SUT-MA-049 | CRD-0601-0434-241218-017_FM_SD_Card_Inserted |
|  | SUT-MA-050 | CRD-0601-0434-241218-018_BattInstOpTask |
| SDS-MA-036 | SUT-MA-046 | CRD-0601-0434-241218-014_AED_mainStart |
| SDS-MA-037 | SUT-MA-046 | CRD-0601-0434-241218-014_AED_mainStart |
| SDS-MA-038 | SUT-MA-046 | CRD-0601-0434-241218-014_AED_mainStart |
| SDS-MA-039 | SUT-MA-051 | CRD-0601-0434-241218-019_EmergencyMsgTask |
| SDS-MA-040 | SUT-MA-023 | CRD-0601-0434-241217-022_PadsCheckTask |
| SDS-MA-041 | SUT-MA-052 | CRD-0601-0434-241218-020_ChargeProcTask |
| SDS-MA-042 | SUT-MA-053 | CRD-0601-0434-241218-021_AEDProcTask |
| SDS-MA-043 | SUT-MA-053 | CRD-0601-0434-241218-021_AEDProcTask |
|  | SUT-MA-054 | CRD-0601-0434-241218-022_ShockProcTask |
|  | SUT-MA-055 | CRD-0601-0434-241218-023_NoShockProcTask |
| SDS-MA-044 | SUT-MA-035 | CRD-0601-0434-241218-003_CPRProcTask |
|  | SUT-MA-056 | CRD-0601-0434-241218-024_Mute_CPR |
|  | SUT-MA-057 | CRD-0601-0434-241218-025_Mute_CPR_END |
|  | SUT-MA-058 | CRD-0601-0434-241218-026_Voice_CPR |
|  | SUT-MA-059 | CRD-0601-0434-241218-027_Voice_CPR_END |
| SDS-MA-045 | SUT-MA-060 | CRD-0601-0434-241218-028_ErrCodePrompt |
|  | SUT-MA-061 | CRD-0601-0434-241218-029_playSoftwareVersion |
|  | SUT-MA-062 | CRD-0601-0434-241218-030_playHardwareVersion |
|  | SUT-MA-063 | CRD-0601-0434-241218-031_InfoReqOpTask |
| SDS-MA-046 | SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
| SDS-MA-047 | SUT-MA-013 | CRD-0601-0434-241217-013_Dev_Config_Data_UpDate |
|  | SUT-MA-063 | CRD-0601-0434-241218-031_InfoReqOpTask |
|  | SUT-MA-065 | CRD-0601-0434-241218-033_CPRSettingPasswordInput |
| SDS-MA-048 | SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
|  | SUT-MA-066 | CRD-0601-0434-241218-034_CommSelectTask |
|  | SUT-MA-067 | CRD-0601-0434-241218-035_SD_Copy_Patient_Data |
| SDS-MA-049 | SUT-MA-047 | CRD-0601-0434-241218-015_FactoryModeChk |
| SDS-MA-050 | SUT-MA-068 | CRD-0601-0434-241218-036_Init_Config_and_Operation_Data |
| SDS-MA-051 | SUT-MA-069 | CRD-0601-0434-241218-037_VOC_SD_to_QSPI |
| SDS-MA-052 | SUT-MA-068 | CRD-0601-0434-241218-036_Init_Config_and_Operation_Data |
| SDS-MA-053 | SUT-MA-061 | CRD-0601-0434-241218-029_playSoftwareVersion |
|  | SUT-MA-062 | CRD-0601-0434-241218-030_playHardwareVersion |
|  | SUT-MA-070 | CRD-0601-0434-241218-038_FactoryPromptTask |
|  | SIT-DS-035 | [CRD-0601-0434-241205-005] SIT-DS-035_Main ARM Menu Item Change Test |
|  | SIT-DS-036 | [CRD-0601-0434-241205-006] SIT-DS-036_Main Menu Entry Test for ARM Submenu_Parent Menu Navigation |
| SDS-MA-054 | SUT-MA-019 | CRD-0601-0434-241217-019_PadsProcTask |
|  | SUT-MA-047 | CRD-0601-0434-241218-015_FactoryModeChk |
|  | SUT-MA-071 | CRD-0601-0434-241218-039_FactoryDefibCtrlTest |
|  | SUT-MA-072 | CRD-0601-0434-241218-040_FactoryAdultSelectPrompt |
|  | SUT-MA-073 | CRD-0601-0434-241218-041_FactoryPediatricSelectPrompt |
|  | SUT-MA-110 | CRD-0601-0434-241219-020_ExtADC_Get |
|  | SIT-DS-006 | [CRD-0601-0569-241202-001] SIT-DS-006_Main ARM_ Norflash Memory Data Retention Test |
|  | SIT-DS-073 | [CRD-0601-0434-241211-004] SIT-DS-073_Main ARM_ i-Button Action Recognition Test |
|  | SIT-DS-074 | [CRD-0601-0434-241211-005] SIT-DS-074_Main ARM_ Shock Button Operational Recognition Test |
|  | SIT-DS-062 | [CRD-0601-0434-241210-003] SIT-DS-062_Main ARM_ i-Button Red LED Output Control Test |
|  | SIT-DS-071 | [CRD-0601-0434-241211-002] SIT-DS-071_Main ARM_ Audio Output Test |
|  | SIT-DS-024 | [CRD-0601-0434-241204-004] SIT-DS-024_Main ARM_Button Combination Recognition Test |
|  | SIT-DS-036 | [CRD-0601-0434-241205-006] SIT-DS-036_Main Menu Entry Test for ARM Submenu_Parent Menu Navigation |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
|  | SIT-DS-069 | [CRD-0601-0569-241210-005] SIT-DS-069_Main ARM_ Beep Sound Output Control Test |
|  | SIT-DS-035 | [CRD-0601-0434-241205-005] SIT-DS-035_Main ARM Menu Item Change Test |
|  | SIT-DS-024 | [CRD-0601-0434-241204-004] SIT-DS-024_Main ARM_Button Combination Recognition Test |
|  | SIT-DS-037 | [CRD-0601-0569-241205-001] SIT-DS-037_Main ARM Device Reference Value Data Validation Test |
|  | SIT-DS-038 | [CRD-0601-0569-241205-002] SIT-DS-038_Main ARM_Multilingual Configured Value Determination Test |
| SDS-MA-055 | SUT-MA-074 | CRD-0601-0434-241218-042_SDCardSysUpdate |
|  | SIT-DS-033 | [CRD-0601-0434-241205-003] SIT-DS-033_Main ARM_SD Card Data Integrity Test |
|  | SIT-DS-034 | [CRD-0601-0434-241205-004] SIT-DS-034_Main ARM_SD Card Data Upload Validation Test |
| SDS-MA-056 | SUT-MA-022 | CRD-0601-0434-241217-022_PadsCheckTask |
|  | SIT-DS-039 | [CRD-0601-0569-241205-003] SIT-DS-039_Main ARM Defibrillation Pads Impedance ADC Measurement Value Voice Output Test |
| SDS-MA-057 | SUT-MA-075 | CRD-0601-0434-241218-043_FM_Detail_Short_Set |
|  | SUT-MA-076 | CRD-0601-0434-241218-044_FM_Language_set |
|  | SIT-DS-035 | [CRD-0601-0434-241205-005] SIT-DS-035_Main ARM Menu Item Change Test |
|  | SIT-DS-036 | [CRD-0601-0434-241205-006] SIT-DS-036_Main Menu Entry Test for ARM Submenu_Parent Menu Navigation |
|  | SIT-DS-038 | [CRD-0601-0569-241205-002] SIT-DS-038_Main ARM_Multilingual Configured Value Determination Test |
| SDS-MA-058 | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
|  | SUT-MA-027 | CRD-0601-0434-241217-027_StopSystemWhenBatteryDepleted |
|  | SUT-MA-077 | CRD-0601-0434-241218-045_PwrOnSelfTest |
|  | SUT-MA-078 | CRD-0601-0434-241218-046_POST_DisarmTest |
|  | SIT-DS-017 | [CRD-0601-0569-241203-001] SIT-DS-017_Main ARM_Battery Remaining Capacity Calculation Test |
|  | SIT-DS-018 | [CRD-0601-0569-241203-002] SIT-DS-018_Main ARM_Battery Voltage ADC Measurement Value Determination Test |
|  | SIT-DS-019 | [CRD-0601-0569-241203-003] SIT-DS-019_Main ARM_Battery Level Determination Test |
|  | SIT-DS-037 | [CRD-0601-0569-241205-001] SIT-DS-037_Main ARM Device Reference Value Data Validation Test |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
| SDS-MA-059 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
| SDS-MA-059.1 | SUT-MA-016 | CRD-0601-0434-241217-016_SystemErrorMsg |
| SDS-MA-059.2 | SUT-MA-025 | CRD-0601-0434-241217-025_BatteryLevelCalculation_withCounter |
| SDS-MA-059.3 | SUT-MA-079 | CRD-0601-0434-241218-047_StateLCDDisplay |
| SDS-MA-059.4 | SIT-DS-011 | [CRD-0601-0434-241203-001] SIT-DS-011_Main ARM Device Inspection Results Error Detection Test |
| SDS-MA-060 | SUT-MA-020 | CRD-0601-0434-241217-020_Pads_EEPROM_Use_Task |
|  | SUT-MA-021 | CRD-0601-0434-241217-021_Set_Pads_Exp_Flag |
|  | SUT-MA-022 | CRD-0601-0434-241217-022_PadsCheckTask |
|  | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
|  | SUT-MA-027 | CRD-0601-0434-241217-027_StopSystemWhenBatteryDepleted |
|  | SUT-MA-080 | CRD-0601-0434-241218-048_PSTOpTask |
|  | SUT-MA-081 | CRD-0601-0434-241218-049_ReadTemperature |
|  | SIT-DS-012 | [CRD-0601-0434-241203-002] SIT-DS-012_Main ARM Defibrillation Pads Connector ADC Measurement Value Determination Test |
|  | SIT-DS-013 | [CRD-0601-0434-241203-003] SIT-DS-013_Main ARM_ Defibrillation Pads EEPROM Usage History Data Validation Test |
|  | SIT-DS-014 | [CRD-0601-0434-241203-004] SIT-DS-014_Main ARM_ Defibrillation Pads EEPROM Expiration Date Validation Test |
|  | SIT-DS-015 | [CRD-0601-0434-241203-005] SIT-DS-015_Main ARM_ Defibrillation Pads Impedance ADC Measurement Value Determination Test |
|  | SIT-DS-017 | [CRD-0601-0569-241203-001] SIT-DS-017_Main ARM_Battery Remaining Capacity Calculation Test |
|  | SIT-DS-018 | [CRD-0601-0569-241203-002] SIT-DS-018_Main ARM_Battery Voltage ADC Measurement Value Determination Test |
|  | SIT-DS-019 | [CRD-0601-0569-241203-003] SIT-DS-019_Main ARM_Battery Level Determination Test |
|  | SIT-DS-037 | [CRD-0601-0569-241205-001] SIT-DS-037_Main ARM Device Reference Value Data Validation Test |
|  | SIT-DS-040 | [CRD-0601-0569-241205-004] SIT-DS-040_Main ARM_Device Shelf Life Data Categorization Test |
|  | SIT-DS-126 | [CRD-0601-0569-241218-001] SIT-DS-124_Main ARM_Patient Setup Mode Change Test |
| SDS-MA-061 | SUT-MA-080 | CRD-0601-0434-241218-048_PSTOpTask |
|  | SUT-MA-082 | CRD-0601-0434-241218-050_ChargeMsgCtrl |
|  | SUT-MA-083 | CRD-0601-0434-241218-051_DefiMsgCtrl |
|  | SUT-MA-084 | CRD-0601-0434-241218-052_DefiEnergyCtrl |
|  | SUT-MA-085 | CRD-0601-0434-241218-053_batteryStatusfromVoltage |
|  | SIT-DS-040 | [CRD-0601-0569-241205-004] SIT-DS-040_Main ARM_Device Shelf Life Data Categorization Test |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
| SDS-MA-062 | SUT-MA-080 | CRD-0601-0434-241218-048_PSTOpTask |
|  | SUT-MA-086 | CRD-0601-0434-241218-054_ChargeDisarmTest |
|  | SIT-DS-040 | [CRD-0601-0569-241205-004] SIT-DS-040_Main ARM_Device Shelf Life Data Categorization Test |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
| SDS-MA-063 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
|  | SUT-MA-016 | CRD-0601-0434-241217-016_SystemErrorMsg |
|  | SUT-MA-025 | CRD-0601-0434-241217-025_BatteryLevelCalculation_withCounter |
|  | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
|  | SUT-MA-079 | CRD-0601-0434-241218-047_StateLCDDisplay |
|  | SUT-MA-080 | CRD-0601-0434-241218-048_PSTOpTask |
|  | SUT-MA-087 | CRD-0601-0434-241218-055_PIC_LED_COLOR_SET |
|  | SUT-MA-088 | CRD-0601-0434-241218-056_AlarmBuzzerThreeTimes |
|  | SUT-MA-089 | CRD-0601-0434-241218-057_RedInfLEDTask |
|  | SIT-DS-005 | [CRD-0601-0434-241202-005] SIT-DS-005_Main ARM_PST Duration Setting Test Procedure |
|  | SIT-DS-006 | [CRD-0601-0569-241202-001] SIT-DS-006_Main ARM_ Norflash Memory Data Retention Test |
|  | SIT-DS-007 | [CRD-0601-0569-241202-002] SIT-DS-007_Main ARM_EEPROM Data Storage Test |
|  | SIT-DS-010 | [CRD-0601-0569-241202-005] SIT-DS-010_Main ARM Error Code Classification Test |
|  | SIT-DS-011 | [CRD-0601-0434-241203-001] SIT-DS-011_Main ARM Device Inspection Results Error Detection Test |
|  | SIT-DS-060 | [CRD-0601-0434-241210-001] SIT-DS-060_Main ARM_ Main ARM I2C Data Transmission Test |
|  | SIT-DS-062 | [CRD-0601-0434-241210-003] SIT-DS-062_Main ARM_ i-Button Red LED Output Control Test |
|  | SIT-DS-069 | [CRD-0601-0569-241210-005] SIT-DS-069_Main ARM_ Beep Sound Output Control Test |
|  | SIT-DS-085 | [CRD-0601-0569-241212-001] SIT-DS-085_Main PIC_ Main PIC I2C Data Transmission Test |
| SDS-MA-064 | SUT-MA-018 | CRD-0601-0434-241217-018_PadsConnectChkMsg |
|  | SUT-MA-020 | CRD-0601-0434-241217-020_Pads_EEPROM_Use_Task |
|  | SUT-MA-021 | CRD-0601-0434-241217-021_Set_Pads_Exp_Flag |
|  | SUT-MA-022 | CRD-0601-0434-241217-022_PadsCheckTask |
|  | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
|  | SUT-MA-027 | CRD-0601-0434-241217-027_StopSystemWhenBatteryDepleted |
|  | SUT-MA-050 | CRD-0601-0434-241218-018_BattInstOpTask |
|  | SUT-MA-082 | CRD-0601-0434-241218-050_ChargeMsgCtrl |
|  | SUT-MA-083 | CRD-0601-0434-241218-051_DefiMsgCtrl |
|  | SUT-MA-084 | CRD-0601-0434-241218-052_DefiEnergyCtrl |
|  | SUT-MA-086 | CRD-0601-0434-241218-054_ChargeDisarmTest |
|  | SUT-MA-090 | CRD-0601-0434-241218-058_ShockRampTask |
|  | SUT-MA-091 | CRD-0601-0434-241219-001_InfoBlueRampTask |
|  | SUT-MA-092 | CRD-0601-0434-241219-002_PushInfoMsgTask |
|  | SUT-MA-093 | CRD-0601-0434-241219-003_ShkKeyTestMsgTask |
|  | SIT-DS-002 | [CRD-0601-0434-241202-002] SIT-DS-002_Main ARM_Internal Memory Data Initialization Test |
|  | SIT-DS-012 | [CRD-0601-0434-241203-002] SIT-DS-012_Main ARM Defibrillation Pads Connector ADC Measurement Value Determination Test |
|  | SIT-DS-013 | [CRD-0601-0434-241203-003] SIT-DS-013_Main ARM_ Defibrillation Pads EEPROM Usage History Data Validation Test |
|  | SIT-DS-014 | [CRD-0601-0434-241203-004] SIT-DS-014_Main ARM_ Defibrillation Pads EEPROM Expiration Date Validation Test |
|  | SIT-DS-015 | [CRD-0601-0434-241203-005] SIT-DS-015_Main ARM_ Defibrillation Pads Impedance ADC Measurement Value Determination Test |
|  | SIT-DS-017 | [CRD-0601-0569-241203-001] SIT-DS-017_Main ARM_Battery Remaining Capacity Calculation Test |
|  | SIT-DS-018 | [CRD-0601-0569-241203-002] SIT-DS-018_Main ARM_Battery Voltage ADC Measurement Value Determination Test |
|  | SIT-DS-019 | [CRD-0601-0569-241203-003] SIT-DS-019_Main ARM_Battery Level Determination Test |
|  | SIT-DS-024 | [CRD-0601-0434-241204-004] SIT-DS-024_Main ARM_Button Combination Recognition Test |
|  | SIT-DS-032 | [CRD-0601-0434-241205-002] SIT-DS-032_Main ARM_NorFlash Memory Data Initialization Test |
|  | SIT-DS-037 | [CRD-0601-0569-241205-001] SIT-DS-037_Main ARM Device Reference Value Data Validation Test |
|  | SIT-DS-041 | [CRD-0601-0434-241206-001] SIT-DS-041_Main ARM_ BIT Voice Data Discrimination Test |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
| SDS-MA-065 | SUT-MA-007 | CRD-0601-0434-241217-007_Dev_Operation_Data_UpDate |
|  | SUT-MA-015 | CRD-0601-0434-241217-015_SystemMonitorTask |
|  | SUT-MA-016 | CRD-0601-0434-241217-016_SystemErrorMsg |
|  | SUT-MA-025 | CRD-0601-0434-241217-025_BatteryLevelCalculation_withCounter |
|  | SUT-MA-026 | CRD-0601-0434-241217-026_LowBattIndicTask |
|  | SUT-MA-079 | CRD-0601-0434-241218-047_StateLCDDisplay |
|  | SUT-MA-087 | CRD-0601-0434-241218-055_PIC_LED_COLOR_SET |
|  | SIT-DS-006 | [CRD-0601-0569-241202-001] SIT-DS-006_Main ARM_ Norflash Memory Data Retention Test |
|  | SIT-DS-007 | [CRD-0601-0569-241202-002] SIT-DS-007_Main ARM_EEPROM Data Storage Test |
|  | SIT-DS-011 | [CRD-0601-0434-241203-001] SIT-DS-011_Main ARM Device Inspection Results Error Detection Test |
|  | SIT-DS-041 | [CRD-0601-0434-241206-001] SIT-DS-041_Main ARM_ BIT Voice Data Discrimination Test |
|  | SIT-DS-060 | [CRD-0601-0434-241210-001] SIT-DS-060_Main ARM_ Main ARM I2C Data Transmission Test |
|  | SIT-DS-062 | [CRD-0601-0434-241210-003] SIT-DS-062_Main ARM_ i-Button Red LED Output Control Test |
|  | SIT-DS-071 | [CRD-0601-0434-241211-002] SIT-DS-071_Main ARM_ Audio Output Test |
|  | SIT-DS-073 | [CRD-0601-0434-241211-004] SIT-DS-073_Main ARM_ i-Button Action Recognition Test |
| SDS-MA-066 | SUT-MA-094 | CRD-0601-0434-241219-004_HVWrCmdProc |
|  | SIT-DS-042 | [CRD-0601-0434-241206-002] SIT-DS-042_Main ARM_Test for Determining Electrical Energy Charge Quantity |
|  | SIT-DS-043 | [CRD-0601-0434-241206-003] SIT-DS-043_Main ARM_Electrical Energy Charging Determination Test |
|  | SIT-DS-044 | [CRD-0601-0434-241206-004] SIT-DS-044_Main ARM_Electrical Energy Discharge Determination Test |
| SDS-MA-067 | SUT-MA-094 | CRD-0601-0434-241219-004_HVWrCmdProc |
|  | SUT-MA-095 | CRD-0601-0434-241219-005_HVRdCmdProc |
|  | SIT-DS-045 | [CRD-0601-0569-241206-001] SIT-DS-045_Main ARM_ Therapy BD Receiver Waiting Time Determination Test |
|  | SIT-DS-046 | [CRD-0601-0569-241206-002] SIT-DS-046_Main ARM_ Therapy BD Operation Status Determination Test |
|  | SIT-DS-047 | [CRD-0601-0569-241206-003] SIT-DS-047_Main ARM_ High Voltage CAP ADC Measurement Value Determination Test |
| SDS-MA-068 | SUT-MA-032 | CRD-0601-0434-241217-032_ECGAnalyzeTask |
|  | SIT-DS-115 | [CRD-0601-0569-241217-001] SIT-DS-115_Main ARM_ ECG Waveform Sensitivity Analysis Test |
| SDS-MA-069 | SUT-MA-032 | CRD-0601-0434-241217-032_ECGAnalyzeTask |
|  | SIT-DS-116 | [CRD-0601-0569-241217-002] SIT-DS-116_Main ARM_ ECG Waveform Specificity Analysis Test |
| SDS-MA-070 | SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
|  | SUT-MA-102 | CRD-0601-0434-241219-012_HVRdCmd |
|  | SUT-MA-103 | CRD-0601-0434-241219-013_HVWrCmd |
|  | SUT-MA-104 | CRD-0601-0434-241219-014_PwrCtrlGet |
|  | SUT-MA-105 | CRD-0601-0434-241219-015_PwrCtrlPut |
|  | SUT-MA-107 | CRD-0601-0434-241219-017_CommHostMode |
|  | SIT-DS-121 | [CRD-0601-0569-241218-002] SIT-DS-121_Main ARM_ CPR Impedance ADC Value Measurement Test |
| SDS-MA-071 | SUT-MA-096 | CRD-0601-0434-241219-006_IntADC_Init |
|  | SUT-MA-097 | CRD-0601-0434-241219-007_IntADCPwrOn |
|  | SUT-MA-098 | CRD-0601-0434-241219-008_IntADCPwrOff |
|  | SUT-MA-099 | CRD-0601-0434-241219-009_IntADCGet |
|  | SIT-DS-048 | [CRD-0601-0569-241206-004] SIT-DS-048_Main ARM_ Defibrillation Pads Connector ADC Measurement Test |
| SDS-MA-072 | SUT-MA-096 | CRD-0601-0434-241219-006_IntADC_Init |
|  | SUT-MA-097 | CRD-0601-0434-241219-007_IntADCPwrOn |
|  | SUT-MA-098 | CRD-0601-0434-241219-008_IntADCPwrOff |
|  | SUT-MA-099 | CRD-0601-0434-241219-009_IntADCGet |
|  | SIT-DS-049 | [CRD-0601-0569-241206-005] SIT-DS-049_Main ARM_Battery Voltage ADC Measurement Test Procedure |
| SDS-MA-073 | SUT-MA-096 | CRD-0601-0434-241219-006_IntADC_Init |
|  | SUT-MA-097 | CRD-0601-0434-241219-007_IntADCPwrOn |
|  | SUT-MA-098 | CRD-0601-0434-241219-008_IntADCPwrOff |
|  | SUT-MA-099 | CRD-0601-0434-241219-009_IntADCGet |
|  | SIT-DS-050 | [CRD-0601-0434-241209-001] SIT-DS-050_Main ARM_Microphone ADC Measurement Test |
| SDS-MA-074 | SUT-MA-096 | CRD-0601-0434-241219-006_IntADC_Init |
|  | SUT-MA-097 | CRD-0601-0434-241219-007_IntADCPwrOn |
|  | SUT-MA-098 | CRD-0601-0434-241219-008_IntADCPwrOff |
|  | SUT-MA-099 | CRD-0601-0434-241219-009_IntADCGet |
|  | SIT-DS-051 | [CRD-0601-0434-241209-002] SIT-DS-051_Main ARM_Temperature Sensor ADC Measurement Test Procedure |
| SDS-MA-075 | SUT-MA-100 | CRD-0601-0434-241219-010_SD_Card_Update |
|  | SIT-DS-052 | [CRD-0601-0434-241209-003] SIT-DS-052_Main ARM_SD Card Insertion Status Recognition Test |
|  | SIT-DS-053 | [CRD-0601-0434-241209-004] SIT-DS-053_Main ARM_SD Card Format Recognition Test |
|  | SIT-DS-055 | [CRD-0601-0434-241209-006] SIT-DS-055_Main ARM_SD Card Data Read Test |
| SDS-MA-076 | SUT-MA-101 | CRD-0601-0434-241219-011_SD_Copy_Patient_Data_Load |
|  | SIT-DS-052 | [CRD-0601-0434-241209-003] SIT-DS-052_Main ARM_SD Card Insertion Status Recognition Test |
|  | SIT-DS-053 | [CRD-0601-0434-241209-004] SIT-DS-053_Main ARM_SD Card Format Recognition Test |
|  | SIT-DS-054 | [CRD-0601-0434-241209-005] SIT-DS-054_Main ARM_SD Card Data Write Test |
| SDS-MA-077 | SUT-MA-102 | CRD-0601-0434-241219-012_HVRdCmd |
|  | SIT-DS-059 | [CRD-0601-0569-241209-004] SIT-DS-059_Main ARM_Main ARM UART Data Reception Test |
| SDS-MA-078 | SUT-MA-103 | CRD-0601-0434-241219-013_HVWrCmd |
|  | SIT-DS-058 | [CRD-0601-0569-241209-003] SIT-DS-058_Main ARM_Main ARM UART Data Transmission Test |
| SDS-MA-079 | SUT-MA-104 | CRD-0601-0434-241219-014_PwrCtrlGet |
|  | SIT-DS-061 | [CRD-0601-0434-241210-002] SIT-DS-061_Main ARM_ Main ARM I2C Data Reception Test |
| SDS-MA-080 | SUT-MA-105 | CRD-0601-0434-241219-015_PwrCtrlPut |
|  | SIT-DS-060 | [CRD-0601-0434-241210-001] SIT-DS-060_Main ARM_ Main ARM I2C Data Transmission Test |
| SDS-MA-081 | SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
|  | SUT-MA-106 | CRD-0601-0434-241219-016_CommBurstTxMode |
|  | SIT-DS-056 | [CRD-0601-0569-241209-001] SIT-DS-056_Main ARM_ IrDA Data Transmission Test |
| SDS-MA-082 | SUT-MA-064 | CRD-0601-0434-241218-032_CommProtocolTask |
|  | SUT-MA-107 | CRD-0601-0434-241219-017_CommHostMode |
|  | SIT-DS-057 | [CRD-0601-0569-241209-002] SIT-DS-057_Main ARM_ IrDA Data Reception Test |
| SDS-MA-083 | SUT-MA-108 | CRD-0601-0434-241219-018_ExtADC_Init |
|  | SUT-MA-109 | CRD-0601-0434-241219-019_ExtADC_Enable |
|  | SUT-MA-110 | CRD-0601-0434-241219-020_ExtADC_Get |
|  | SIT-DS-099 | [CRD-0601-0569-241213-004] SIT-DS-099_Main ARM Defibrillation Pads Impedance ADC Measurement Test |
| SDS-MA-084 | SUT-MA-108 | CRD-0601-0434-241219-018_ExtADC_Init |
|  | SUT-MA-109 | CRD-0601-0434-241219-019_ExtADC_Enable |
|  | SUT-MA-110 | CRD-0601-0434-241219-020_ExtADC_Get |
|  | SIT-DS-100 | [CRD-0601-0434-241216-001] SIT-DS-100_Main ARM_ECG Signal ADC Measurement Test |
| SDS-MA-085 | SUT-MA-111 | CRD-0601-0434-241219-021_ExtRTC_Get |
|  | SUT-MA-112 | CRD-0601-0434-241219-022_ExtRTC_Put |
|  | SIT-DS-060 | [CRD-0601-0434-241210-001] SIT-DS-060_Main ARM_ Main ARM I2C Data Transmission Test |
|  | SIT-DS-061 | [CRD-0601-0434-241210-002] SIT-DS-061_Main ARM_ Main ARM I2C Data Reception Test |
| SDS-MA-086 | SUT-MA-113 | CRD-0601-0434-241219-023_LED_InfoRED |
|  | SUT-MA-114 | CRD-0601-0434-241219-024_LED_InfoBLUE |
|  | SUT-MA-115 | CRD-0601-0434-241219-025_LED_ShockRamp |
|  | SUT-MA-116 | CRD-0601-0434-241219-026_LED_Power_Switch_Ramp |
|  | SUT-MA-117 | CRD-0601-0434-241219-027_LED_Patient_Contact_Ban |
|  | SUT-MA-118 | CRD-0601-0434-241219-028_LED_Pads_Indication_Ramp |
|  | SUT-MA-119 | CRD-0601-0434-241219-029_LED_Patient_Contact_Good |
|  | SIT-DS-062 | [CRD-0601-0434-241210-003] SIT-DS-062_Main ARM_ i-Button Red LED Output Control Test |
|  | SIT-DS-063 | [CRD-0601-0434-241210-004] SIT-DS-063_Main ARM_ i-Button Blue LED Output Control Test |
|  | SIT-DS-064 | [CRD-0601-0434-241210-005] SIT-DS-064_Main ARM_ Shock Button LED Output Control Test |
|  | SIT-DS-065 | [CRD-0601-0569-241210-001] SIT-DS-065_Main ARM_ Power Button LED Output Control Test |
|  | SIT-DS-066 | [CRD-0601-0569-241210-002] SIT-DS-066_Main ARM_ Do-Not Touch Patient Indicator Output Control Test |
|  | SIT-DS-067 | [CRD-0601-0569-241210-003] SIT-DS-067_Main ARM_ Pad Position Indicators Output Control Test |
|  | SIT-DS-068 | [CRD-0601-0569-241210-004] SIT-DS-068_Main ARM_ CPR Detection Indicator Output Control Test |
| SDS-MA-087 | SUT-MA-120 | CRD-0601-0434-241219-030_Buzzer_Timer_Init |
|  | SUT-MA-121 | CRD-0601-0434-241219-031_Buzzer_ON |
|  | SUT-MA-122 | CRD-0601-0434-241219-032_Buzzer_OFF |
|  | SIT-DS-069 | [CRD-0601-0569-241210-005] SIT-DS-069_Main ARM_ Beep Sound Output Control Test |
| SDS-MA-088 | SUT-MA-123 | CRD-0601-0434-241219-033_Audio_VolumeSet |
|  | SUT-MA-124 | CRD-0601-0434-241219-034_AudioPlayCtrl |
|  | SIT-DS-070 | [CRD-0601-0434-241211-001] SIT-DS-070_Main ARM_ Audio Volume Adjustment Test |
|  | SIT-DS-071 | [CRD-0601-0434-241211-002] SIT-DS-071_Main ARM_ Audio Output Test |
| SDS-MA-089 | SUT-MA-044 | CRD-0601-0434-241218-012_InfantPadsMsgTask |
|  | SUT-MA-045 | CRD-0601-0434-241218-013_ModeChangeTask |
|  | SUT-MA-125 | CRD-0601-0434-241219-035_KEYPAD_InfKeyChkOn |
|  | SUT-MA-126 | CRD-0601-0434-241219-036_KEYPAD_ShkKeyChkOn |
|  | SUT-MA-127 | CRD-0601-0434-241219-037_KEYPAD_PwrKeyChkOn |
|  | SIT-DS-072 | [CRD-0601-0434-241211-003] SIT-DS-072_Main ARM_power button Recognition Test of Operation |
|  | SIT-DS-073 | [CRD-0601-0434-241211-004] SIT-DS-073_Main ARM_ i-Button Action Recognition Test |
|  | SIT-DS-074 | [CRD-0601-0434-241211-005] SIT-DS-074_Main ARM_ Shock Button Operational Recognition Test |
|  | SIT-DS-075 | [CRD-0601-0434-241211-006] SIT-DS-075_Main ARM_AdultPediatric Selection Switch Operational Recognition Test |
| SDS-MA-090 | SUT-MA-081 | CRD-0601-0434-241218-049_ReadTemperature |
|  | SIT-DS-126 | [CRD-0601-0569-241218-001] SIT-DS-124_Main ARM_Patient Setup Mode Change Test |
| SDS-MP-001 | SUT-MP-001 | CRD-0601-0434-241216-001_RunModeTask |
|  | SIT-DS-076 | [CRD-0601-0569-241211-001] SIT-DS-076_Main ARM Error Detection Test Using Watchdog Timer |
| SDS-MP-002 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-077 | [CRD-0601-0569-241211-002] SIT-DS-077_Main PIC_ Wakeup State Configuration Test |
| SDS-MP-003 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-077 | [CRD-0601-0569-241211-002] SIT-DS-077_Main PIC_ Wakeup State Configuration Test |
| SDS-MP-004 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-077 | [CRD-0601-0569-241211-002] SIT-DS-077_Main PIC_ Wakeup State Configuration Test |
| SDS-MP-005 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-077 | [CRD-0601-0569-241211-002] SIT-DS-077_Main PIC_ Wakeup State Configuration Test |
| SDS-MP-006 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-077 | [CRD-0601-0569-241211-002] SIT-DS-077_Main PIC_ Wakeup State Configuration Test |
| SDS-MP-007 | SUT-MP-003 | CRD-0601-0434-241216-003_LoopWithSleepModeForever |
| SDS-MP-008 | SUT-MP-002 | CRD-0601-0434-241216-002_ChgRunMode |
|  | SIT-DS-079 | [CRD-0601-0569-241211-004] SIT-DS-079_Main PIC_ Main ARM Power Activation Test |
| SDS-MP-009 | SUT-MP-004 | CRD-0601-0434-241216-004_ChgSleepMode |
|  | SIT-DS-080 | [CRD-0601-0434-241212-001] SIT-DS-080_Main PIC_ Test for Deactivating Main ARM Power |
| SDS-MP-010 | SUT-MP-004 | CRD-0601-0434-241216-004_ChgSleepMode |
|  | SIT-DS-082 | [CRD-0601-0434-241212-003] SIT-DS-082_Main PIC System Clock Control Test |
| SDS-MP-011 | SUT-MP-005 | CRD-0601-0434-241216-005_SleepModeTask |
|  | SIT-DS-123 | [CRD-0601-0569-241218-004] SIT-DS-123_Main ARM CPR Determination of Action Test |
| SDS-MP-012 | SUT-MP-005 | CRD-0601-0434-241216-005_SleepModeTask |
|  | SIT-DS-083 | [CRD-0601-0434-241212-004] SIT-DS-083_Main PIC_ RTC Time Measurement Test |
| SDS-MP-013 | SUT-MP-006 | CRD-0601-0434-241216-006_IICIsr |
|  | SUT-MP-007 | CRD-0601-0434-241216-007_IICEnable |
|  | SUT-MP-008 | CRD-0601-0434-241216-008_IICDisable |
|  | SIT-DS-084 | [CRD-0601-0434-241212-005] SIT-DS-084_Main PIC_ Main PIC I2C Data Reception Test |
| SDS-MP-014 | SUT-MP-006 | CRD-0601-0434-241216-006_IICIsr |
|  | SUT-MP-007 | CRD-0601-0434-241216-007_IICEnable |
|  | SUT-MP-008 | CRD-0601-0434-241216-008_IICDisable |
|  | SIT-DS-085 | [CRD-0601-0569-241212-001] SIT-DS-085_Main PIC_ Main PIC I2C Data Transmission Test |
| SDS-MP-015 | SUT-MP-006 | CRD-0601-0434-241216-006_IICIsr |
|  | SIT-DS-084 | [CRD-0601-0434-241212-005] SIT-DS-084_Main PIC_ Main PIC I2C Data Reception Test |
|  | SIT-DS-085 | [CRD-0601-0569-241212-001] SIT-DS-085_Main PIC_ Main PIC I2C Data Transmission Test |
| SDS-MP-016 | SUT-MP-009 | CRD-0601-0434-241216-009_ds2431_init |
|  | SIT-DS-086 | [CRD-0601-0569-241212-002] SIT-DS-086_Main PIC 1-Wire Communication Initialization Test |
| SDS-MP-017 | SUT-MP-010 | CRD-0601-0434-241216-010_OW_read_byte |
|  | SUT-MP-011 | CRD-0601-0434-241216-011_ds2431_read_eeprom |
|  | SIT-DS-088 | [CRD-0601-0569-241212-004] SIT-DS-088_Main PIC_EEPROM Read Test Procedure |
| SDS-MP-018 | SUT-MP-012 | CRD-0601-0434-241216-012_OW_write_byte |
|  | SUT-MP-013 | CRD-0601-0434-241216-013_ds2431_write_eeprom |
|  | SIT-DS-087 | [CRD-0601-0569-241212-003] SIT-DS-087_Main PIC_EEPROM Write Test Procedure |
| SDS-MP-019 | SUT-MP-010 | CRD-0601-0434-241216-010_OW_read_byte |
|  | SUT-MP-011 | CRD-0601-0434-241216-011_ds2431_read_eeprom |
|  | SIT-DS-088 | [CRD-0601-0569-241212-004] SIT-DS-088_Main PIC_EEPROM Read Test Procedure |
| SDS-MP-020 | SUT-MP-012 | CRD-0601-0434-241216-012_OW_write_byte |
|  | SUT-MP-013 | CRD-0601-0434-241216-013_ds2431_write_eeprom |
|  | SIT-DS-087 | [CRD-0601-0569-241212-003] SIT-DS-087_Main PIC_EEPROM Write Test Procedure |
| SDS-MP-021 | SUT-MP-013 | CRD-0601-0434-241216-013_ds2431_write_eeprom |
|  | SIT-DS-007 | [CRD-0601-0569-241202-002] SIT-DS-007_Main ARM_EEPROM Data Storage Test |
|  | SIT-DS-120 | [CRD-0601-0434-241218-001] SIT-DS-120_Main ARM Defibrillation Pads Information Data |
| SDS-MP-022 | SUT-MP-013 | CRD-0601-0434-241216-013_ds2431_write_eeprom |
|  | SIT-DS-007 | [CRD-0601-0569-241202-002] SIT-DS-007_Main ARM_EEPROM Data Storage Test |
|  | SIT-DS-119 | [CRD-0601-0569-241217-005] SIT-DS-119_Main PIC_Battery Information Data |
| SDS-MP-023 | SUT-MP-014 | CRD-0601-0434-241216-014_PadsOWReset |
|  | SUT-MP-015 | CRD-0601-0434-241216-015_LCDUpdateDuringSleep |
|  | SUT-MP-016 | CRD-0601-0434-241216-016_LED_UPDATE |
|  | SIT-DS-081 | [CRD-0601-0434-241212-002] SIT-DS-081_Main PIC Defibrillation Pads Connector Status Determination Test |
| SDS-MP-024 | SUT-MP-016 | CRD-0601-0434-241216-016_LED_UPDATE |
|  | SIT-DS-089 | [CRD-0601-0569-241212-005] SIT-DS-089_Main PIC_ Status Indicator (Red Light) Output Control Test |
| SDS-MP-025 | SUT-MP-016 | CRD-0601-0434-241216-016_LED_UPDATE |
|  | SIT-DS-090 | [CRD-0601-0434-241213-001] SIT-DS-090_Main PIC_ Status Indicator (Green Light) Output Control Test |
| SDS-MP-026 | SUT-MP-005 | CRD-0601-0434-241216-005_SleepModeTask |
|  | SIT-DS-091 | [CRD-0601-0434-241213-002] SIT-DS-091_Main PIC_ Power Button Recognition Test |
| SDS-MP-027 | SUT-MP-005 | CRD-0601-0434-241216-005_SleepModeTask |
|  | SIT-DS-092 | [CRD-0601-0434-241213-003] SIT-DS-092_Main PIC_ i-Button Recognition Test |
| SDS-MP-028 | SUT-MP-017 | CRD-0601-0434-241216-017_Set_LCD_Low_When_True |
|  | SUT-MP-018 | CRD-0601-0434-241216-018_Set_LCD_High_When_True |
|  | SIT-DS-093 | [CRD-0601-0434-241213-004] SIT-DS-093_Main PIC_ Status LCD (Device Operation) Output Test (O |
| SDS-MP-029 | SUT-MP-017 | CRD-0601-0434-241216-017_Set_LCD_Low_When_True |
|  | SUT-MP-018 | CRD-0601-0434-241216-018_Set_LCD_High_When_True |
|  | SIT-DS-097 | [CRD-0601-0569-241213-002] SIT-DS-097_Main PIC_ Status LCD (Pads Status) High Output Control Test |
|  | SIT-DS-098 | [CRD-0601-0569-241213-003] SIT-DS-098_Main PIC_ Status LCD (Pads Status) Low Output Control Test |
| SDS-MP-030 | SUT-MP-017 | CRD-0601-0434-241216-017_Set_LCD_Low_When_True |
|  | SUT-MP-018 | CRD-0601-0434-241216-018_Set_LCD_High_When_True |
|  | SIT-DS-094 | [CRD-0601-0434-241213-005] SIT-DS-094_Main PIC_ Status LCD (Battery Level Indicator) High Output Control Test |
|  | SIT-DS-095 | [CRD-0601-0434-241213-006] SIT-DS-095_Main PIC_ Status LCD (Battery Level Indicator) Middle Area Output Control Test |
|  | SIT-DS-096 | [CRD-0601-0569-241213-001] SIT-DS-096_Main PIC_ Status LCD (Battery Level Indicator) Low Level Output Control Test |
| SDS-TP-001 | SUT-TP-001 | CRD-0601-0434-241213-001_Charge |
|  | SIT-DS-101 | [CRD-0601-0434-241216-002] SIT-DS-101_Therapy PIC_ Therapy PIC Control Command Judgment Test |
| SDS-TP-002 | SUT-TP-001 | CRD-0601-0434-241213-001_Charge |
|  | SIT-DS-101 | [CRD-0601-0434-241216-002] SIT-DS-101_Therapy PIC_ Therapy PIC Control Command Judgment Test |
| SDS-TP-003 | SUT-TP-001 | CRD-0601-0434-241213-001_Charge |
|  | SIT-DS-101 | [CRD-0601-0434-241216-002] SIT-DS-101_Therapy PIC_ Therapy PIC Control Command Judgment Test |
|  | SIT-DS-104 | [CRD-0601-0434-241216-005] SIT-DS-104_Therapy PIC_Test for Disruption of Electrical Energy Charging Stop |
| SDS-TP-004 | SUT-TP-002 | CRD-0601-0434-241213-002_BiteDischgOff |
|  | SUT-TP-003 | CRD-0601-0434-241213-003_Discharge |
|  | SIT-DS-101 | [CRD-0601-0434-241216-002] SIT-DS-101_Therapy PIC_ Therapy PIC Control Command Judgment Test |
|  | SIT-DS-105 | [CRD-0601-0569-241216-001] SIT-DS-105_Therapy PIC_External Discharge Test for Electrical Energy |
| SDS-TP-005 | SUT-TP-004 | CRD-0601-0434-241213-004_Disarm |
|  | SIT-DS-101 | [CRD-0601-0434-241216-002] SIT-DS-101_Therapy PIC_ Therapy PIC Control Command Judgment Test |
|  | SIT-DS-106 | [CRD-0601-0569-241216-002] SIT-DS-106_Therapy PIC_Electrical Energy Internal Discharge Test |
| SDS-TP-006 | SUT-TP-005 | CRD-0601-0434-241213-005_FirstPhaseDischgOn |
|  | SUT-TP-006 | CRD-0601-0434-241213-006_FirstPhaseDischgOff |
|  | SUT-TP-007 | CRD-0601-0434-241213-007_SecondPhaseDischgOn |
|  | SUT-TP-008 | CRD-0601-0434-241213-008_SecondPhaseDischgOff |
|  | SIT-DS-107 | [CRD-0601-0569-241216-003] SIT-DS-107_Therapy PIC_Electrical Energy Output Control Test |
| SDS-TP-007 | SUT-TP-009 | CRD-0601-0434-241213-009_AdcInit |
|  | SUT-TP-010 | CRD-0601-0434-241213-010_AdcSetChannel |
|  | SUT-TP-011 | CRD-0601-0434-241213-011_AdcRead |
|  | SIT-DS-108 | [CRD-0601-0569-241216-004] SIT-DS-108_Therapy PIC_ High Voltage CAP ADC Measurement Test |
| SDS-TP-008 | SUT-TP-012 | CRD-0601-0434-241213-012_main |
|  | SUT-TP-013 | CRD-0601-0434-241213-013_RxISR |
|  | SIT-DS-110 | [CRD-0601-0434-241217-001] SIT-DS-110_TherapyPIC_ Therapy PIC UART Data Reception Test |
| SDS-TP-009 | SUT-TP-014 | CRD-0601-0434-241213-014_TxISR |
|  | SIT-DS-109 | [CRD-0601-0569-241216-005] SIT-DS-109_Therapy PIC_ Therapy PIC UART Data Transmission Test |